## TOC & Survey Qs
SWRCB. (2021). Laboratory Survey Responses: November 2021. State Water Resources Control Board.
Table of Contents
Aluminum
Antimony
Arsenic
Barium
Beryllium
Boron
Cadmium
Chromium (total)
Copper
Iron
Lead
Manganese
Mercury
Nickel
Selenium
Silver
Thallium
Vanadium
Zinc
List of Questions
1a. Analytical method(s) used (Please use dropdown menu to select primary analytical method. If more than one method is used, please use the second "Response" column.)
1b. What percent of sample analyses for this analyte is conducted using the analytical method selected? (ONLY ANSWER IF using more than one analytical method for this analyte)
2a. Minimum Reporting Limit (μg/L) (Please type in value)
2b. Range of typical spiked sample recoveries for your analysis of this analyte (e.g., ±30, ±50%) with 99% confidence. (Please type in range)
2c. MRL with 99% confidence for recovery within 70% to 130%? (Please type in value)
3a. Can you detect down to the target DLR?
3b. Lowest calibration point (μg/L) (Please type in value)
4a. MDL(b) (μg/L) (Based on Method Blanks) (Please type in value)
4b. MDL determination procedure for 4a.?
5a. MDL(s) (μg/L) (Based on Spiked Samples) (Please type in value)
5b. MDL determination procedure for 5a.?
6a. Cost of Analysis using Method based on Current Regulation ($) (Please type in value)
6b. Is the sample analysis performed based on a single element? Or suite of elements?
6c. Is there a pricing discount available for contracted/regular clients? (If yes, please type in reduced sample analysis cost. If no, please skip to next question.)
6d. At a DLR [see below], what percentage increase in analysis cost would you expect, if any? [10 times lower than the current DLR] for Aluminum, Antimony, Arsenic, Barium, Beryllium, Copper, Lead, Mercury, Nickel, Selenium, and Titanium [100 times lower than the current notification level] for Boron and Vanadium [100 times lower than the current DLR] for Cadmium and Chromium (total) [100 times lower than the current secondary MCL] for Iron, Manganese, Silver, and Zinc
6e. If DLR is lowered to the level specified in 6d., would you need to purchase new instrumentation/ equipment?
6f. Approximate cost of new instrumentation/equipment needed? (Please type in value; leave blank if not applicable)
6g. If DLR is lowered would your lab maintain ELAP Accreditation?

## Aluminum
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to current DLR of 50 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 5 ug/L | 6e. Equipment Needed for DLR of 5 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | NA | 10 | ±30% | 10 | Yes | 10 | 3.3 | As described in 2017 Methods Update Rule | 5.226 | As described in method | $35 (Estimated for single metal) | Suite | NA - Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | Not sure | Yes | Not sure | Most likely, however, will need to address issues with common contamination of blanks around 5 ppb
| Unknown | 50 | ±30% | 50 | Yes | 0 | 3.9 | 2016 EPA MDL Procedure | 19 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | Unknown | Unknown. | Unknown | No
| 30% | 50 | ±30% | 50 | Yes | 10,000 (1-point calibration) | 41.5 | As described in 2017 Methods Update Rule | 19.4 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | NA | New higher acid purity and teflon digestion cup | NA | No
| 80%+ | 50 | ±50% | 75 | Yes | 50 | 12.2 | As described in 2017 Methods Update Rule | 16.5 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None (Price will increase over time) | NA | NA | NA | No
| NA | 50 | ±30% | 200 | Yes | 50,000 (1-point calibration) | 27 | As described in 2017 Methods Update Rule | 39 | As described in 2017 Methods Update Rule | $25 (Discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 80%+ | 50 | 90-110% | 80-120% | Yes | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | No
| NA | 50 | ±10% | NA | Yes | 5000 | NA | NA | 14.2 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 50 | ±15% | NA | Yes | (1-point calibration) | 12 | As described in method | NA | NA | 22 | Single | NA | NA | No (As long as our instrument is suitable to see that low (ICP-AES), then we would be ok.) | NA | Yes
| 10% | 50 | 85-115% | Uknown | Yes | 10 | 20 | As described in 2017 Methods Update Rule | 20 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Uknown | No
EPA 200.8 | 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.662 | As described in 2017 Methods Update Rule | 0.119 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | greater than 30% | No | NA | Yes
| 10% | 3 | ±30% | 3.3 | Yes | 3 | 0.307 | As described in 2017 Methods Update Rule | 0.639 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite | None (Price will increase over time) | 40% | No | NA | Yes
| NA | 4 | ±30% | NA | Yes | 4 | 2.9 | As described in 2017 Methods Update Rule | 1.8 | As described in 2017 Methods Update Rule | 200 | Suite | NA | 20% | No | 0 | Yes
| 80%+ | 5 | ±30% | 5 | Yes | 1 | 1 | As described in 2017 Methods Update Rule | 0.8 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| 80%+ | 10 | 85-115% | 10 | Yes | 8 | 8 | As described in 2017 Methods Update Rule | 4 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| NA | 10 | ±30% | 10 | Yes | 10 | 0.396 | As described in 2017 Methods Update Rule | 0.458 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Yes | No Idea | Maybe
| NA | 10 | ±30% | 10 | Yes | 10 | 0.0728 (undigested blank) | As described in method | 0.467 (undigested sample) | As described in method | We perform the analysis in-house | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 15 | ±30% | 15 | Yes | 10 | 1 | As described in 2017 Methods Update Rule | 5 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| 80%+ | 20 | ±25% | 20 | Yes | 10 | 4.8 | As described in 2017 Methods Update Rule | 6.8 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| 75% | 20 | ±30% | 20 | Yes | 0.5 | 4.43 | As described in 2017 Methods Update Rule | 2.52 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | greater than 30% | Yes | 25000 | Yes
| 80%+ | 20 | 70-134% | 20 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 20 | 4.2 | As described in 2017 Methods Update Rule | 1.13 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| Unknown | 50 | ±30% | 50 | Yes | 0 | NA | NA | 0.92 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 50 | NA | NA | Yes | 50 | 18 | As described in 2017 Methods Update Rule | 16 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 5 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 1 | 4 | As described in 2017 Methods Update Rule | 4 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Antimony
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to MRL of 1 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.6 ug/L | 6e. Equipment Needed for DLR of 0.6 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.5 | NA | 6 | ±15% | NA | Yes | (1-point calibration) | 1 | As described in method | NA | NA | 22 | Single | NA | NA | NA | NA | Yes (As long as our instrument is suitable to see that low (ICP-AES), then we would be ok)
EPA 200.8 | NA | 0.1 | ±70% | 0.1 | Yes | 0.1 | 0.355 | As described in 2017 Methods Update Rule | 0.021 | As described in 2017 Methods Update Rule | 20 | Single | NA | No idea | Yes | No Idea | Maybe
| NA | 0.4 | ±30% | 0.44 | Yes | 0.4 | 0.0182 | As described in 2017 Methods Update Rule | 0.0408 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| NA | 0.5 | ±30% | 1 | Yes | 0.1 | 0.056 | As described in 2017 Methods Update Rule | 0.089 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 ug/L | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | less than 5% | No | NA | Yes
| NA | 0.5 | ±30% | NA | Yes | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.06 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | No | 100000 | Yes
| NA | 0.5 | 85-115% | Unknown | Yes | 0.4 | 0.03 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| 80%+ | 1 | ±25% | 1 | Yes | 1 | 0.14 | As described in 2017 Methods Update Rule | 0.033 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| 80%+ | 1 | ±30% | 1 | Yes | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.3 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 0 | NA | NA | 0.03 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.06 | As described in 2017 Methods Update Rule | 0.024 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.387 | As described in 2017 Methods Update Rule | 0.269 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 1 | 85-125% | 1 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 1 | 0.148 | As described in 2017 Methods Update Rule | 0.064 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | NA | NA | Yes | 1 | 0.61 | As described in 2017 Methods Update Rule | 0.21 | As described in 2017 Methods Update Rule | NA | NA | NA | NA | NA | NA | NA
| NA | 2 | ±30% | 2 | No | 2 | 0.0676 (Undigested Blank) | As described in method | 0.467 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (the cost of an ICP with triple quad MS) | Yes
| NA | 2 | ±30% | 200 | Yes | 50 (1-point calibration) | 0.048 | As described in 2017 Methods Update Rule | 0.059 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 6 | +/- 10% | NA | Yes | 1 | NA | NA | 2.937 | As described in method | 20 | Single | It depends on # of samples | greater than 30% | Yes | NA | Yes
| NA | 6 | NA | NA | No | 1 | 1.3 | As described in 2017 Methods Update Rule | 0.73 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 0.3 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.1 | 0.04 | As described in 2017 Methods Update Rule | 0.03 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Arsenic
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to MRL of 0.004 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.2 ug/L | 6e. Equipment Needed for DLR of 0.2 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.5 (run only for non-regulatory samples - homeowners) | NA | 2 | ±15% | NA | Yes | (1-point calibration) | 0.8 | As described in method | NA | NA | 22 | Single | NA | NA | No (As long as our instrument is suitable to see that low (ICP-AES), then we would be ok) | NA | No (200.5 is not accepted for As anymore)
EPA 200.8 | NA | 0.1 | ±70% | 0.1 | No | 0.1 | 0.005 | As described in 2017 Methods Update Rule | 0.006 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Yes | No Idea | Maybe
| NA | 0.4 | ±30% | 0.5 | No | 0.1 | 0.023 | As described in 2017 Methods Update Rule | 0.074 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 | No | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 20% | No | NA | Yes
| NA | 0.5 | ±30% | 0.5 | No | 0.5 | 0.0316 (Undigested Blank) | As described in method | 0.0457 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 0.5 | 85-115% | Unknown | No | 0.4 | 0.3 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| NA | 0.8 | ±30% | 0.88 | No | 0.8 | 0.103 | As described in 2017 Methods Update Rule | 0.19 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (may request single analyte within the suite) | None | greater than 30% (Specialty gas, new instrument purchase and labor costs will result in costs much greater than 30%) | No | NA | No
| 80%+ | 1 | ±25% | 1 | No | 1 | 0.15 | As described in 2017 Methods Update Rule | 0.09 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| 80%+ | 1 | ±30% | 1 | No | 1 | 1 | As described in 2017 Methods Update Rule | 0.8 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | 30% | Yes | 100000 | No
| NA | 1 | ±50% | 10 | No | 1 | 0.032 | As described in 2017 Methods Update Rule | 0.046 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 1 | 82-130% | 1 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 1 | 0.077 | As described in 2017 Methods Update Rule | 0.149 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±10% | NA | NA | 1 | NA | NA | 0.198 | As described in 2017 Methods Update Rule | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 1 | ±30% | NA | No | 1 | 0.3 | As described in 2017 Methods Update Rule | 0.5 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| NA | 2 | ±30% | 2 | No | 0 | NA | NA | 0.34 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 25% | No | NA | Yes
| NA | 2 | ±30% | 2 | No | 1 | 0.188 | As described in 2017 Methods Update Rule | 0.385 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | Would require the use of collision cell technology which is not approved for drinking water | Not sure | NA | Not sure
| NA | 2 | ±30% | 200 | No | 50 (1-point calibration) | 0.68 | As described in 2017 Methods Update Rule | 0.83 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 2 | ±20% | NA | No | 2 | 0.38 | As described in 2017 Methods Update Rule | 0.37 | As described in 2017 Methods Update Rule | 200 | Suite | NA | greater than 30% | Yes | NA | No
| NA | 2 | NA | NA | No | 1 | 0.56 | As described in 2017 Methods Update Rule | 0.66 | As described in 2017 Methods Update Rule | 15 | suite | NA | 10% | No | NA | Yes
| NA | 0.5 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | No | 0.5 | 0.3 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | Yes (0.2 ug/L) | Unknown | Yes
SM 3113B | NA | 2 | ±30% | 2 | No | 2 | 0.5 | As described in 2017 Methods Update Rule | 0.6 | As described in 2017 Methods Update Rule | 40 | Single | NA | greater than 30% | Yes | 240000 | Yes

## Barium
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to current DLR of 100 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 10 ug/L | 6e. Equipment Needed for DLR of 10 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.5 | 50% | 100 | ±15% | NA | Yes | (1-point calibration) | NA | NA | NA | NA | 22 | Single | NA | NA | NA | NA | Yes
EPA 200.7 | Unknown | 1 | ±30% | 1 | Yes | 0 | 0.078 | 2016 EPA MDL Procedure | 0.16 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 25% | No | NA | Probably
| 10% | 2 | ±30% | 2 | Yes | 1,000 (1-point calibration) | 0.395 | As described in 2017 Methods Update Rule | 0.433 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 10% | 10 | 85-115% | 10 | Yes | 10 | 0.4 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | No | NA | Yes
| 80%+ | 50 | ±30% | 55 | Yes | 50 | 0.408 | As described in 2017 Methods Update Rule | 0.391 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 50% | 50 | ±30% | 200 | Yes | 10000 | 1.3 | As described in 2017 Methods Update Rule | 2.4 | As described in 2017 Methods Update Rule | $25 (Prices discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes
| NA | 100 | 90-110% | 80-120% | Yes | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | Yes
| NA | 100 | ±15% | NA | Yes | 1000 | NA | NA | 11.5 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| 50% | 100 | ±15% | NA | Yes | (1-point calibration) | NA | NA | NA | NA | 22 | Single | NA | NA | NA | NA | Yes
EPA 200.8 | 20% | 0.2 | ±50% | 0.3 | Yes | 0.2 | 0.0298 | As described in 2017 Methods Update Rule | 0.0296 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 80%+ | 0.5 | ±30% | 0.05 | Yes | 1 | 0.03 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| 80%+ | 0.5 | 85-115% | 0.5 | Yes | 0.5 | 0.4 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 1 | ±25% | 1 | Yes | 0.2 | 0.044 | As described in 2017 Methods Update Rule | 0.095 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | No | 140000 | Yes
| NA | 1 | ±70% | 1 | Yes | 1 | 0.154 | As described in 2017 Methods Update Rule | 0.249 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Yes | No idea | maybe
| Unknown | 1 | ±30% | 1 | Yes | 0 | NA | NA | 0.021 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.032 | As described in 2017 Methods Update Rule | 0.098 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients | less than 5% | No | NA | Yes
| 80%+ | 1 | ±30% | 1 | Yes | 0.1 | 0.07 | As described in 2017 Methods Update Rule | 0.142 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.011 | As described in 2017 Methods Update Rule | 0.041 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 4 | NA | NA | Yes | 4 | 2.4 | As described in 2017 Methods Update Rule | 1.97 | As described in 2017 Methods Update Rule | 200 | Suite | NA | 5% | No | NA | Yes
| NA | 5 | ±30% | 5 | Yes | 5 | 0.0184 (Undigested Blank) | As described in method | 0.155 (Undigested Sample) | As described in method | NA | Suite | NA | NA | No | NA | Yes
| 50% | 5 | ±30% | 200 | Yes | 50 (1-point calibration) | 0.13 | As described in 2017 Methods Update Rule | 0.13 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 5 | ±30% | NA | Yes | 1 | 0.2 | As described in 2017 Methods Update Rule | 0.5 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | less than 5% | No | NA | Yes
| NA | 100 | NA | NA | Yes | 50 | 6.7 | As described in 2017 Methods Update Rule | 6.7 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 0.1 (PQL) | 70-130 (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.1 | 0.02 | As described in 2017 Methods Update Rule | 0.02 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes
| 80%+ | NA | 2 | 2 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 2 | 0.112 | As described in 2017 Methods Update Rule | 0.108 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes

## Beryllium
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to current DLR of 1 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.1 ug/L | 6e. Equipment Needed for DLR of 0.1 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.5 | 50% | 1 | ±15% | NA | NA | (1-point calibration) | 0.1 | As described in method | NA | NA | 22 | Single | NA | NA | NA | NA | Yes
EPA 200.7 | Unknown | 1 | ±30% | 1 | Yes | 0 | 0.23 | 2016 EPA MDL Procedure | 0.02 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 25% | No | NA | Probably
| 75% | 1 | ±30% | 1.1 | Yes | 1 | 0.243 | As described in 2017 Methods Update Rule | 0.243 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (may request single analyte within the suite) | None | less than 5% | No | NA | No
| 50% | 1 | ±15% | NA | NA | (1-point calibration) | 1 | As described in method | NA | NA | 22 | Single | NA | NA | NA | NA | Yes
| 10% | 1 | 85-115% | Unknown | Yes | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.3 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Uknown | No
EPA 200.8 | NA | 0.1 | ±70% | 0.01 | Yes | 0.1 | 0.005 | As described in 2017 Methods Update Rule | 0.005 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| NA | 0.1 | ±30% | 0.1 | Yes | 0.1 | 0.023 | As described in 2017 Methods Update Rule | 0.062 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 40% | 0.1 | ±30% | 0.11 | Yes | 0.1 | 0.0095 | As described in 2017 Methods Update Rule | 0.112 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (may request single analyte within the suite) | None | greater than 30% (Specialty gas, new instrument purchase and labor costs will result in costs much greater than 30%) | No | NA | No
| 80%+ | 0.1 | 85-115% | Unknown | Yes | 0.1 | 0.09 | As described in 2017 Methods Update Rule | 0.06 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| 80%+ | 0.5 | ±30% | 0.05 | Yes | 1 | 0.06 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | 10% | No | NA | Yes
| NA | 0.5 | ±30% | 0.5 | Yes | 0.5 | 0.00983 (Undigested Blank) | As described in method | 0.111 (Undigested Sample) | As described in method | NA | Single | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 0.5 | NA | NA | Yes | 0.5 | 0.145 | As described in 2017 Methods Update Rule | 0.12 | As described in 2017 Methods Update Rule | 200 | Suite | NA | greater than 30% | Yes | NA | No
| NA | 0.5 | ±30% | NA | Yes | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.01 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | 25% | No | 25000 | Yes
| 80%+ | 1 | ±25% | 1 | Yes | 1 | 0.089 | As described in 2017 Methods Update Rule | 0.019 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| Unknown | 1 | ±30% | 1 | Yes | 0 | NA | NA | 0.058 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.008 | As described in 2017 Methods Update Rule | 0.012 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.099 | As described in 2017 Methods Update Rule | 0.116 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±30% | 200 | Yes | 50 (1-point calibration) | 0.018 | As described in 2017 Methods Update Rule | 0.21 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 80%+ | 1 | 85-122% | 1 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 1 | 0.011 | As described in 2017 Methods Update Rule | 0.11 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±10% | NA | Yes | 1 | NA | NA | 0.09 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | Yes | NA | Yes
| NA | 1 | NA | NA | Yes | 1 | 0.38 | As described in 2017 Methods Update Rule | 0.6 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 20% | No | NA | Yes
| NA | 0.1 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.1 | 0.02 | As described in 2017 Methods Update Rule | 0.05 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Boron
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to the target MRL? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 10 ug/L | 6e. Equipment Needed for DLR of 10 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | NA | 5 | ±30% | 5 | NA | 0 | 0 | 2016 EPA MDL Procedure | 0.33 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| 80%+ | 10 | ±25% | 10 | NA | 2-point calibration: 0 and 10,000 | 0.044 | As described in 2017 Methods Update Rule | 0.057 | As described in 2017 Methods Update Rule | $175 (includes Anions, TDS, EC) | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 10 | ±30% | 10 | Yes | 1000 | 2.65 | As described in 2017 Methods Update Rule | 2.69 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| NA | 50 | ±30% | 55 | NA | 50 | 9.11 | As described in 2017 Methods Update Rule | 9.6 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | 30% (Specialty gas, new instrument purchase and labor costs will result in costs much greater than 30%) | No | NA | No
| 80%+ | 50 | 81-119% | 50 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 50 | 19.2 | As described in 2017 Methods Update Rule | 33.9 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 100 | ±30% | 100 | NA | 100 | 3.4 | As described in 2017 Methods Update Rule | 3.8 | As described in method | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | Contamination of blanks around 10 ppb | No | NA | Not sure
| NA | 100 | ±30% | 200 | NA | 10,000 (1-point calibration) | 18 | As described in 2017 Methods Update Rule | 19 | As described in 2017 Methods Update Rule | $25 (Discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 100 | ±50% | 200 | Yes | 100 | 84 | As described in 2017 Methods Update Rule | 71 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 100 | 90-110% | 80-90% | Yes | 100 | 100 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | 5% | Yes | 150000 | No
| NA | 100 | ±15% | NA | NA | 5000 | NA | NA | 70.4 | As described in 2017 Methods Update Rule | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 100 | NA | NA | NA | 100 | 6.2 | As described in 2017 Methods Update Rule | 29 | As described in 2017 Methods Update Rule | 15 | Suite | NA | 10% | No | NA | Yes
| 50% | 100 | 85-115% | Unknown | NA | 100 | 90 | As described in 2017 Methods Update Rule | 30 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Unknown | No
EPA 200.8 | 80%+ | 0.5 | ±15% | 0.5 | NA | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | less than 5% | No | NA | Yes
| NA | 20 | ±30% | 20 | Yes | 20 | 7.887 (Undigested Blank) | As described in method | 6.409 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (This is the cost of an ICP with triple quad MS) | Yes
| 80%+ | 50 | ±30% | 20 | NA | 20 | 20 | As described in 2017 Methods Update Rule | 30 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | greater than 30% | Yes | 100000 | No
| 50% | 50 | 85-115% | Unknown | NA | 50 | 11 | As described in 2017 Methods Update Rule | 8 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| NA | 100 | ±70% | 100 | NA | 100 | 3 | As described in 2017 Methods Update Rule | 16 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Maybe | No Idea | Maybe
| NA | 5 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | NA | 1 | 0.7 | As described in 2017 Methods Update Rule | 4 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Cadmium
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to DLR of 0.04 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.01 ug/L | 6e. Equipment Needed for DLR of 0.01 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | Unknown | 1 | ±30% | 1 | No | 0 | 0.19 | 2016 EPA MDL Procedure | 0.21 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 100% | Unknown | Unknown | No
| 40% | 1 | ±30% | 1.1 | No | 1 | 0.315 | As described in 2017 Methods Update Rule | 0.369 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | Unknown if any technology can meet the proposed DLR of 0.01 ug/L | NA | NA | No
EPA 200.8 | NA | 0.1 | ±70% | 0.1 | No | 0.1 | 0.028 | As described in 2017 Methods Update Rule | 0.013 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No ldea | Maybe
| NA | 0.1 | ±30% | 0.1 | No | 0.1 | 0.00557 (Undigested Blank) | As described in method | 0.0046 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (This is the cost of an ICP with triple quad MS) | Yes
| 75% | 0.1 | ±30% | 0.11 | Yes | 0.1 | 0.0132 | As described in 2017 Methods Update Rule | 0.0144 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | Unknown if any technology can meet the proposed DLR of 0.01 ug/L | No | NA | No
| NA | 0.1 | 85-115% | Unknown | Yes | 0.1 | 0.04 | As described in 2017 Methods Update Rule | 0.03 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| 80%+ | 0.2 | ±25% | 0.2 | No | 0.2 | 0.021 | As described in 2017 Methods Update Rule | 0.043 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 0.2 | ±30% | 0.2 | No | 0.1 | 0.028 | As described in 2017 Methods Update Rule | 0.042 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | greater than 30% | No | NA | Yes
| NA | 0.3 | ±30% | NA | No | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| 80%+ | 0.5 | ±30% | 0.1 | No | 1 | 0.15 | As described in 2017 Methods Update Rule | 0.3 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | greater than 30% | Yes | 100000 | No
| 80%+ | 0.5 | ±15% | 0.5 | No | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 20% | No | NA | Yes
| 80%+ | 0.5 | 85-118% | 0.5 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 0.5 | 0.028 | As described in 2017 Methods Update Rule | 0.038 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 0.5 | NA | NA | No | 0.5 | 0.12 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 200 | Suite | NA | greater than 30% | Yes | NA | No
| Unknown | 1 | ±30% | 1 | Yes | 0 | NA | NA | 0.011 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 10% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.004 | As described in 2017 Methods Update Rule | 0.026 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | Not sure | Not sure | NA | Not sure
| NA | 1 | ±50% | 10 | No | 1 | 0.144 | As described in 2017 Methods Update Rule | 0.158 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | No
| NA | 1 | ±30% | 200 | No | 50 (1-point calibration) | 0.031 | As described in 2017 Methods Update Rule | 0.099 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 1 | ±10% | NA | NA | 1 | NA | NA | 0.13 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 1 | NA | NA | No | 1 | 0.78 | As described in 2017 Methods Update Rule | 0.53 | As described in 2017 Methods Update Rule | 15 | Suite | NA | 10% | No | NA | Yes
| NA | 0.1 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.1 | 0.001 | As described in 2017 Methods Update Rule | 0.02 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | Yes (0.01 ug/L) | Unknown | Yes

## Chromium (total)
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to DLR of 0.05 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.1 ug/L | 6e. Equipment Needed for DLR of 0.1 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | Unknown | 5 | ±30% | 5 | No | 0 | 0.34 | 2016 EPA MDL Procedure | 0.44 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 25% | No | NA | Probably
| 75% | 10 | ±30% | 11 | No | 10 | 1.65 | As described in 2017 Methods Update Rule | 1.73 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | Unknown if any technology can meet the proposed DLR of 0.01 ug/L | No | NA | No
| NA | 10 | 90-110% | 80-120% | No | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | NA
| 10% | 10 | 85-115% | Unknown | No | 10 | 2 | As described in 2017 Methods Update Rule | 2 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Unknown | No
EPA 200.8 | NA | 0.2 | ±30% | 0.2 | No | 0.1 | 0.061 | As described in 2017 Methods Update Rule | 0.089 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | greater than 30% | No | NA | Yes
| 40% | 0.4 | ±30% | 0.44 | No | 0.4 | 0.0967 | As described in 2017 Methods Update Rule | 0.111 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | Unknown if any technology can meet the proposed DLR of 0.01 ug/L | Yes | 30000 | No
| 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 20% | No | NA | Yes
| 80%+ | 0.5 | 85-115% | Unknown | No | 0.5 | 0.3 | As described in 2017 Methods Update Rule | 0.3 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| 80%+ | 1 | ±30% | 0.15 | No | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | greater than 30% | Yes | 100000 | No
| 80%+ | 1 | ±25% | 1 | No | 1 | 0.006 | As described in 2017 Methods Update Rule | 0.11 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 120000 | Yes
| NA | 1 | ±70% | 1 | No | 1 | 0.03 | As described in 2017 Methods Update Rule | 0.03 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| NA | 1 | ±30% | 1 | No | 1 | 0.0225 (Undigested Blank) | As described in method | 0.136 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| Unknown | 1 | ±30% | 1 | No | 0 | NA | NA | 0.051 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | No | 1 | 0.068 | As described in 2017 Methods Update Rule | 0.07 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | Not sure | Not sure | Not sure | Not sure
| NA | 1 | ±50% | 10 | No | 1 | 0.191 | As described in 2017 Methods Update Rule | 0.163 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 1 | 85-121% | 1 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 1 | 0.235 | As described in 2017 Methods Update Rule | 0.102 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±30% | NA | No | 1 | 0.2 | As described in 2017 Methods Update Rule | 0.5 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| NA | 5 | NA | NA | No | 5 | 3.3 | As described in 2017 Methods Update Rule | 1.7 | As described in 2017 Methods Update Rule | 200 | Suite | NA | greater than 30% | Yes | NA | No
| NA | 10 | ±30% | 200 | No | 50 (1-point calibration) | 0.95 | As described in 2017 Methods Update Rule | 1.1 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 10 | ±15% | NA | NA | 1 | NA | NA | 0.363 | As described in method | 20 | Single | It depends on the method | greater than 30% | NA | NA | Yes
| NA | 10 | NA | NA | No | 10 | 1.1 | As described in 2017 Methods Update Rule | 0.51 | As described in 2017 Methods Update Rule | 15 | Suite | NA | 10% | No | NA | Yes
| NA | 1 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | No | 0.1 | 1 | As described in 2017 Methods Update Rule | 0.6 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes
SM 3113B | NA | 10 | ±30% | 10 | No | 2 | 0.5 | As described in 2017 Methods Update Rule | 0.6 | As described in 2017 Methods Update Rule | 40 | Single | NA | greater than 30% | Yes | 240000 | Yes

## Copper
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to current DLR of 50 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 5 ug/L | 6e. Equipment Needed for DLR of 5 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | Unknown | 2 | ±30% | 2 | Yes | 0 | 0.23 | 2016 EPA MDL Procedure | 0.27 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| 20% | 10 | ±30% | 10 | Yes | 2,000 (1-point calibration) | 1.55 | As described in 2017 Methods Update Rule | 1.33 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 50% | 20 | ±30% | NA | Yes | 10 | 6 | As described in 2017 Methods Update Rule | 4 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | No | Would switch to 200.8 | No
| 10% | 20 | 85-115% | Unknown | Yes | 20 | 20 | As described in 2017 Methods Update Rule | 8 | NA | 40 | Single | Varies | NA | Yes | Unknown | No
| 50% (304 of 564 by 200.7 in 2021) | 50 | ±30% | 55 | Yes | 50 | 4.56 | As described in 2017 Methods Update Rule | 3.99 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 20% | 50 | ±30% | 200 | Yes | 10000 | 4.9 | As described in 2017 Methods Update Rule | 5.9 | As described in 2017 Methods Update Rule | $25 (Discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes
| NA | 50 | 90-110% | 80-120% | Yes | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | No
| NA | 50 | ±15% | NA | Yes | 5000 | NA | NA | 0.0247 | NA | 20 | Single | It depends proj sample | greater than 30% | NA | NA | Yes
EPA 200.8 | 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 0.5 | ±30% | 0.5 | Yes | 0.1 | 0.19 | As described in 2017 Methods Update Rule | 0.23 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 80%+ | 0.5 | 85-115% | Unknown | Yes | 0.5 | 0.3 | As described in 2017 Methods Update Rule | 0.4 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| 50% (260 of 564 in 2021) | 0.8 | ±50% | 1.2 | Yes | 0.8 | 0.25 | As described in 2017 Methods Update Rule | 0.226 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 80%+ | 1 | ±25% | 1 | Yes | 1 | 0.31 | As described in 2017 Methods Update Rule | 0.94 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 1 | ±70% | 1 | Yes | 1 | 0.034 | As described in 2017 Methods Update Rule | 0.091 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Maybe | No Idea | Maybe
| Unknown | 1 | ±30% | 1 | Yes | 0 | NA | NA | 0.028 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.054 | As described in 2017 Methods Update Rule | 0.072 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.169 | As described in 2017 Methods Update Rule | 0.17 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 2 | 66-129% | 2 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 2 | 0.099 | As described in 2017 Methods Update Rule | 0.092 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 2 | NA | NA | Yes | 2 | 0.6 | As described in 2017 Methods Update Rule | 0.72 | As described in 2017 Methods Update Rule | 200 | Suite | NA | less than 5% | No | NA | Yes
| 50% | 2 | ±30% | NA | Yes | 1 | 0.5 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | 25% | No | 25000 | Yes
| 80%+ | 5 | ±30% | 0.15 | Yes | 1 | 0.2 | As described in 2017 Methods Update Rule | 0.3 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| 80%+ | 5 | ±30% | 200 | Yes | 50 (1-point calibration) | 0.12 | As described in 2017 Methods Update Rule | 0.29 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 75% | 10 | ±30% | 10 | Yes | 10 | 0.0228 (Undigested Blank) | As described in method | 0.525 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 50 | NA | NA | Yes | 50 | 19.2 | As described in 2017 Methods Update Rule | 2.2 | As described in 2017 Methods Update Rule | 15 | suite | NA | less than 5% | No | NA | Yes
| NA | 0.5 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.5 | 0.2 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes
SM 3113B | NA | 50 | ±30% | 50 | Yes | 5 | 0.9 | As described in 2017 Methods Update Rule | 0.6 | As described in 2017 Methods Update Rule | 40 | Single | NA | less than 5% | No | NA | Yes

## Iron
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to the target MRL? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 3 ug/L | 6e. Equipment Needed for DLR of 3 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | NA | 5 | ±50% | 40 | Yes | 5 | 3 | As described in 2017 Methods Update Rule | 1.46 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 10 | ±30% | 10 | NA | 10 | 0.0057 | As described in method | 0.00988 | As described in method | NA | Suite | NA | greater than 30% | Yes (Using EPA 200.7, it will be very challenging to meet this proposed limit) | NA | Yes
| NA | 10 | ±30% | 10 | NA | 10 | 1.8 | As described in 2017 Methods Update Rule | 0.61 | As described in method | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| 80%+ | 10 | 72-129% | 10 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 10 | 0.47 | As described in 2017 Methods Update Rule | 1.4 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 30 | ±25% | 30 | NA | 2-point calibration: 0 and 10,000 | 11 | As described in 2017 Methods Update Rule | 6.6 | As described in 2017 Methods Update Rule | $175 (includes Anions, TDS, EC) | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 30 | ±30% | 30 | NA | 10,000 (1-point calibration) | 2.93 | As described in 2017 Methods Update Rule | 4.95 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | greater than 30% | No | NA | Yes
| NA | 30 | ±30% | 200 | NA | 50,000 (1-point calibration) | 7.3 | As described in 2017 Methods Update Rule | 8.1 | As described in 2017 Methods Update Rule | $25 (Discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 80%+ | 30 | ±30% | NA | NA | 10 | 11 | As described in 2017 Methods Update Rule | 5 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| 80%+ | 50 | ±30% | 50 | NA | 100 | 20 | As described in 2017 Methods Update Rule | 50 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | greater than 30% | Yes | 100000 | No
| NA | 50 | ±30% | 50 | NA | 0 | 7 | 2016 EPA MDL Procedure | 0.59 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 25% | No | NA | Probably
| NA | 50 | ±30% | 55 | NA | 50 | 5.98 | As described in 2017 Methods Update Rule | 10.9 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | Unknown if any technology can meet the proposed DLR of 0.01 ug/L | Yes | 30000 | No
| NA | 50 | 85-115% | Unknown | NA | 50 | 40 | As described in 2017 Methods Update Rule | 20 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Unknown | No
| NA | 100 | 90-110% | 80-120% | NA | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | No
| NA | 100 | ±15% | NA | NA | 5000 | NA | NA | 19.4 | As described in 2017 Methods Update Rule | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 100 | NA | NA | Yes | 100 | 8.4 | As described in 2017 Methods Update Rule | 35 | As described in 2017 Methods Update Rule | 15 | suite | NA | 10% | No | NA | Yes
EPA 200.8 | 10% | 20 | ±30% | NA | NA | 10 | 6 | As described in 2017 Methods Update Rule | 6 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| 80%+ | 50 | ±20% | 50 | NA | 0.5 | NA | NA | 10 | As described in 2017 Methods Update Rule | 12 | Suite | NA | less than 5% | No | NA | Yes
| NA | 100 | ±70% | 100 | NA | 100 | 4 | As described in 2017 Methods Update Rule | 16 | As described in 2017 Methods Update Rule | 20 | Single | NA | No idea | Likely | No Idea | Maybe
| NA | 25 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | NA | 25 | 4 | As described in 2017 Methods Update Rule | 8 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | Yes (3 ug/L) | Unknown | Yes

## Lead
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to DLR of 0.2 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.5 ug/L | 6e. Equipment Needed for DLR of 0.5 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.8 | NA | 0.1 | ±70% | 0.1 | Yes | 0.1 | 0.009 | As described in 2017 Methods Update Rule | 0.008 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| NA | 0.2 | ±30% | 0.2 | Yes | 0.1 | 0.08 | As described in 2017 Methods Update Rule | 0.083 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| NA | 0.25 | 85-115% | 0.25 | Yes | 0.25 | 0.07 | As described in 2017 Methods Update Rule | 0.04 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| NA | 0.4 | ±30% | 0.44 | Yes | 0.4 | 0.0214 | As described in 2017 Methods Update Rule | 0.0298 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 80%+ | 0.5 | ±30% | 0.05 | Yes | 1 | 0.05 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 10% | No | NA | Yes
| 80%+ | 0.5 | ±25% | 0.5 | Yes | 0.2 | 0.005 | As described in 2017 Methods Update Rule | 0.089 | As described in 2017 Methods Update Rule | 170 | Suite | 0.15 | 25% | No | 140000 | Yes
| 80%+ | 0.5 | 0-249% (Lead hits tend to cause wide swings in MS recoveries/spiking. Average recovery is 100.7% over 3 months, but standard deviation is quite large) | 0.5 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 0.5 | 0.031 | As described in 2017 Methods Update Rule | 0.028 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±30% | 1 | No | 1 | 0.02 (Undigested Blank) | As described in method | 0.0314 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 1 | ±30% | 1 | Yes | 0 | NA | NA | 0.014 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.01 | As described in 2017 Methods Update Rule | 0.031 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.122 | As described in 2017 Methods Update Rule | 0.135 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±30% | 200 | No | 50 (1-point calibration) | 0.068 | As described in 2017 Methods Update Rule | 0.14 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 1 | NA | NA | No | 1 | 0.24 | As described in 2017 Methods Update Rule | 0.24 | As described in 2017 Methods Update Rule | 200 | Suite | NA | 20% | No | NA | Yes
| NA | 1 | ±30% | NA | No | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | 25% | No | 25000 | Yes
| NA | 5 | ±10% | NA | NA | 1 | NA | NA | 1.24 | As described in 2017 Methods Update Rule | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 5 | NA | NA | No | 1 | 4.6 | As described in 2017 Methods Update Rule | 0.58 | As described in 2017 Methods Update Rule | 15 | Suite | NA | 10% | No | NA | Yes
| NA | 0.1 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.1 | 0.04 | As described in 2017 Methods Update Rule | 0.03 | As described in 2017 Methods Update Rule | $40/metal | Suite (Both single and suite) | Yes, depends on scope of work. | None | No | NA | Yes
SM 3113B | NA | 5 | ±30% | 5 | No | 5 | 0.8 | As described in 2017 Methods Update Rule | 0.6 | As described in 2017 Methods Update Rule | 40 | Single | NA | greater than 30% | Yes | 240000 | Yes

## Manganese
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to the target MRL? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.5 ug/L | 6e. Equipment Needed for DLR of 0.5 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | 20% | 5 | ±30% | 5 | NA | 2,000 (1-point calibration) | 0.13 | As described in 2017 Methods Update Rule | 0.34 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | greater than 30% | No | NA | Yes
| 80%+ | 10 | ±25% | 10 | NA | 0 (2-point calibration: 0 and 2,000) | 1.9 | As described in 2017 Methods Update Rule | 2.1 | As described in 2017 Methods Update Rule | $175 (includes Anions, TDS, EC) | Suite | 0.15 | 25% | Yes | 140000 | Yes
| NA | 10 | ±30% | 200 | NA | 10,000 (1-point calibration) | 0.8 | As described in 2017 Methods Update Rule | 1.8 | As described in 2017 Methods Update Rule | $25 (Discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 75% | 15 | ±30% | NA | NA | 10 | 4 | As described in 2017 Methods Update Rule | 3 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| 50% | 20 | ±30% | 22 | NA | 20 | 0.234 | As described in 2017 Methods Update Rule | 0.178 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| NA | 20 | 90-110% | 80-120% | NA | 10 | 10 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | No
| NA | 20 | ±15% | NA | NA | 5000 | NA | NA | 0.798 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| 50% | 20 | 85-115% | Unknown | NA | 20 | 2 | As described in 2017 Methods Update Rule | 3 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Unknown | No
| Unknown | 100 | ±30% | 100 | NA | 0 | 0.056 | 2016 EPA MDL Procedure | 0.045 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| 80%+ | NA | ±30% | 10 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
EPA 200.8 | 50% | 0.1 | ±30% | 0.11 | NA | 0.1 | 0.0227 | As described in 2017 Methods Update Rule | 0.0197 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 | NA | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 10% | No | NA | Yes
| 80%+ | 1 | ±30% | 0.05 | NA | 1 | 0.05 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| Unknown | 1 | ±30% | 1 | NA | 0 | NA | NA | 0.035 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| 20% | 1 | ±30% | 1 | NA | 1 | 0.046 | As described in 2017 Methods Update Rule | 0.069 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| 80%+ | 1 | ±30% | 1 | No | 0.1 | 0.11 | As described in 2017 Methods Update Rule | 0.092 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.308 | As described in 2017 Methods Update Rule | 0.255 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±70% | NA | NA | 1 | 0.008 | As described in 2017 Methods Update Rule | 0.061 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| 80%+ | 2 | 67-135% | 2 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 2 | 0.045 | As described in 2017 Methods Update Rule | 0.076 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 2 | NA | NA | NA | 2 | 1.4 | As described in 2017 Methods Update Rule | 0.91 | As described in 2017 Methods Update Rule | 200 | Suite | NA | greater than 30% | Yes | NA | Yes
| 50% | 5 | 85-115% | 5 | NA | 5 | 1 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| 75% | 5 | ±30% | 5 | NA | 5 | 0.0355 (Undigested Blank) | As described in method | 0.4 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| 20% | 5 | ±30% | NA | NA | 1 | 0.2 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| NA | 20 | NA | NA | NA | 10 | 5.3 | As described in 2017 Methods Update Rule | 6.2 | As described in 2017 Methods Update Rule | 15 | Suite | NA | 10% | No | NA | Yes
| NA | 2 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | NA | 0.1 | 0.06 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Mercury
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to current DLR of 1 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.1 ug/L | 6e. Equipment Needed for DLR of 0.1 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.8 | NA | 0.2 | ±30% | 0.2 | Yes | 0.2 | 0.112 (Undigested Blank) | As described in method | 0.0204 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 0.2 | ±30% | 5 | Yes | 0.2 (1-point calibration) | 0.023 | As described in 2017 Methods Update Rule | 0.046 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 80%+ | 0.2 | 81-115% | 0.2 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 0.2 | 0.0016 | As described in 2017 Methods Update Rule | 0.02 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 0.25 | NA | NA | Yes | 0.25 | 0.17 | As described in 2017 Methods Update Rule | 0.05 | As described in 2017 Methods Update Rule | 200 | Suite | NA | 20% | No | NA | Yes
| NA | 0.3 | ±30% | NA | Yes | 0.25 | 0.1 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| 80%+ | 0.5 | ±30% | 0.1 | Yes | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | greater than 30% | Yes | 100000 | No
| NA | 1 | ±70% | 1 | Yes | 1 | 0.135 | As described in 2017 Methods Update Rule | 0.05 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likelly | No Idea | Maybe
| NA | 1 | ±50% | 10 | Yes | 1 | 0.096 | As described in 2017 Methods Update Rule | 0.044 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±10% | NA | Yes | 0.1 | NA | NA | 0.55 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 1 | NA | NA | Yes | 0.5 | 0.2 | As described in 2017 Methods Update Rule | 0.09 | As described in 2017 Methods Update Rule | 15 | Single | NA | 10% | No | NA | Yes
EPA 245.1 | 80%+ | 0.01 | ±25% | 0.01 | Yes | 0.01 | 0.008 | As described in 2017 Methods Update Rule | 0.007 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | No | 140000 | Yes
| NA | 0.02 | 85-115% | Unknown | Yes | 0.02 | 0.02 | As described in 2017 Methods Update Rule | 0.02 | As described in 2017 Methods Update Rule | 80 | Single | Varies | NA | No | NA | Yes
| NA | 0.05 | ±30% | 0.05 | Yes | 0.05 | 0.025 | As described in 2017 Methods Update Rule | 0.029 | As described in 2017 Methods Update Rule | 48 | Single | Yes (varies) | greater than 30% | No | NA | Yes
| NA | 0.05 | ±30% | 0.055 | Yes | 0.05 | 0.0245 | As described in 2017 Methods Update Rule | 0.021 | As described in 2017 Methods Update Rule | 141.9 | Single | None | less than 5% | No | NA | Yes
| NA | 0.2 | 90-110% | 80-120% | Yes | 0.2 | 0.2 | As described in 2017 Methods Update Rule | NA | NA | 60 | Single | Yes | less than 5% | Yes | 150000 | NA
| NA | 1 | ±30% | 1 | Yes | 0.5 | 0.077 | 2016 EPA MDL Procedure | 0.16 | 2016 EPA MDL Procedure | 40 | Single element | It depends | 25% | No | NA | Probably
| NA | 1 | ±30% | 1 | Yes | 0.2 | 0.058 | As described in 2017 Methods Update Rule | 0.055 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | Not sure | Not sure | Not sure | Not sure
| NA | 0.05 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.05 | 0.01 | As described in 2017 Methods Update Rule | 0.004 | As described in 2017 Methods Update Rule | 59 | Single | Yes, depends on scope of work | None | No | NA | Yes

## Nickel
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to current DLR of 10 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 1 ug/L | 6e. Equipment Needed for DLR of 1 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | Unknown | 5 | ±30% | 5 | Yes | 0 | 1.2 | 2016 EPA MDL Procedure | 0.61 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 10% | No | NA | Yes
| 10% | 10 | 85-115% | 10 | Yes | 10 | 3 | As described in 2017 Methods Update Rule | 2 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Unknown | No
| 75% | 10 | ±30% | 11 | Yes | 10 | 1.04 | As described in 2017 Methods Update Rule | 1.08 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | other (explain) | No | NA | No
| NA | 10 | 90-110% | 80-120% | No | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | NA
EPA 200.8 | NA | 0.1 | ±70% | 0.01 | Yes | 0.1 | 0.006 | As described in 2017 Methods Update Rule | 0.015 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 10% | No | NA | Yes
| 80%+ | 0.5 | 85-115% | Unknown | Yes | 0.5 | 0.2 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| 40% | 0.8 | ±30% | 0.88 | Yes | 0.8 | 0.044 | As described in 2017 Methods Update Rule | 0.0303 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite | None | greater than 30% | No | NA | Yes
| 80%+ | 1 | ±30% | 0.05 | Yes | 1 | 0.05 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| 80%+ | 1 | ±25% | 1 | Yes | 1 | 0.74 | As described in 2017 Methods Update Rule | 0.98 | As described in 2017 Methods Update Rule | 170 | Suite | 0.15 | 25% | No | 140000 | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.39 | As described in 2017 Methods Update Rule | 0.479 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.11 | As described in 2017 Methods Update Rule | 0.111 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 2 | ±30% | 2 | Yes | 2 | 0.0167 (Undigested Blank) | As described in method | 0.0921 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| Unknown | 2 | ±30% | 2 | Yes | 0 | NA | NA | 0.12 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends. | 0% | No | NA | Yes
| NA | 2 | ±30% | 2 | Yes | 0.1 | 0.04 | As described in 2017 Methods Update Rule | 0.16 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | greater than 30% | No | NA | Yes
| NA | 2 | NA | NA | Yes | 2 | 0.83 | As described in 2017 Methods Update Rule | 0.43 | As described in 2017 Methods Update Rule | 200 | Suite | NA | 30% | No | NA | Yes
| 80%+ | 5 | 81-114% | 5 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 5 | 0.039 | As described in 2017 Methods Update Rule | 0.33 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 5 | ±30% | NA | Yes | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| NA | 10 | ±30% | 200 | Yes | 50 (1-point calibration) | 0.36 | As described in 2017 Methods Update Rule | 0.38 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 10 | ±10% | NA | Yes | 1 | NA | NA | 0.1033 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 10 | NA | NA | Yes | 10 | 3 | As described in 2017 Methods Update Rule | 3.4 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 0.1 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.1 | 0.07 | As described in 2017 Methods Update Rule | 0.06 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Selenium
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to current DLR of 5 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.5 ug/L | 6e. Equipment Needed for DLR of 0.5 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.8 | NA | 0.4 | ±30% | 0.4 | Yes | 0.1 | 0.054 | As described in 2017 Methods Update Rule | 0.067 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| NA | 0.5 | ±70% | 0.5 | Yes | 0.5 | 0.176 | As described in 2017 Methods Update Rule | 0.046 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| 80%+ | 1 | ±25% | 1 | Yes | 1 | 0.33 | As described in 2017 Methods Update Rule | 0.32 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.032 | As described in 2017 Methods Update Rule | 0.066 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | NA | NA | Yes | 1 | 0.36 | As described in 2017 Methods Update Rule | 0.27 | As described in 2017 Methods Update Rule | 200 | Suite | NA | 10% | No | NA | Yes
| NA | 1 | ±30% | NA | Yes | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.5 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| NA | 2 | ±50% | 3 | Yes | 0.8 | 0.415 | As described in 2017 Methods Update Rule | 0.48 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | greater than 30% | Yes (Purchase of GFAA to meet lower DLR) | 55000 | Yes
| NA | 2 | ±30% | 200 | Yes | 50 (1-point calibration) | 0.56 | As described in 2017 Methods Update Rule | 0.76 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 2 | 85-115% | Unknown | Yes | 2 | 0.5 | As described in 2017 Methods Update Rule | 0.3 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| 80%+ | 2.5 | ±30% | 1.5 | Yes | 5 | 1.5 | As described in 2017 Methods Update Rule | 2 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | greater than 30% | Yes | 100000 | No
| NA | 5 | ±30% | 5 | Yes | 5 | 0.221 (Undigested Blank) | As described in method | 0.252 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 5 | ±30% | 5 | Yes | 0 | NA | NA | 0.39 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 5 | ±30% | 5 | Yes | 1 | 2.022 | As described in 2017 Methods Update Rule | 1.896 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | Would require the use collision cell technology, which is not approved for drinking water | Not sure | NA | Not sure
| 80%+ | 5 | 83-132% | 5 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 5 | 0.101 | As described in 2017 Methods Update Rule | 0.274 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 5 | ±10% | NA | Yes | 1 | NA | NA | 0.1907 | As described in method | 20 | Single | It depends of proj scope | greater than 30% | NA | NA | Yes
| NA | 5 | NA | NA | Yes | 1 | 0.53 | As described in 2017 Methods Update Rule | 0.85 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| 80%+ | 20 | ±15% | 20 | Yes | 0.5 | NA | NA | 4 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 10% | No | NA | Yes
| NA | 2 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.5 | 2 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | Yes (0.5 ug/L) | Unknown | Yes
SM 3113B | NA | 5 | ±30% | 5 | Yes | 2 | 0.9 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 40 | Single | NA | greater than 30% | Yes | 240000 | Yes

## Silver
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to the target MRL? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 1 ug/L | 6e. Equipment Needed for DLR of 1 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | Unknown | 5 | ±30% | 5 | NA | 0 | 0.31 | 2016 EPA MDL Procedure | 1.6 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 10% | No | NA | Yes
| 10% | 10 | 85-115% | 10 | NA | 10 | 3 | As described in 2017 Methods Update Rule | 2 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | Yes | Unknown | No
| 80%+ | 10 | ±50% | 15 | NA | 10 | 3.65 | As described in 2017 Methods Update Rule | 8.31 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | other (explain) | No | NA | No
| 20% | 10 | ±30% | 100 | NA | 2500 | 6.7 | As described in 2017 Methods Update Rule | 4.8 | As described in 2017 Methods Update Rule | $25 (Discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 10 | 90-110% | 80-120% | NA | 10 | 10 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | No
EPA 200.8 | 20% | 0.1 | ±30% | 0.11 | NA | 0.1 | 0.0184 | As described in 2017 Methods Update Rule | 0.0186 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | greater than 30% | No | NA | Yes
| 80%+ | 0.1 | 85-115% | Unknown | NA | 0.1 | 0.1 | As described in 2017 Methods Update Rule | 0.05 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| NA | 0.2 | ±30% | 0.2 | NA | 0.1 | 0.009 | As described in 2017 Methods Update Rule | 0.03 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 | NA | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 10% | No | NA | Yes
| 80%+ | 0.5 | 80-114% | 0.5 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 0.5 | 0.038 | As described in 2017 Methods Update Rule | 0.038 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 0.5 | NA | NA | NA | 0.5 | 0.11 | As described in 2017 Methods Update Rule | 0.12 | As described in 2017 Methods Update Rule | 200 | Suite | NA | 5% | No | NA | Yes
| 80%+ | 1 | ±25% | 1 | NA | 1 | 0.053 | As described in 2017 Methods Update Rule | 0.11 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 1 | ±50% | 1 | NA | 1 | 0.157 | As described in 2017 Methods Update Rule | 0.051 | As described in 2017 Methods Update Rule | 20 | Single | NA | No dea | Likely | No Idea | Maybe
| NA | 1 | ±30% | 1 | NA | 1 | 0.013 | As described in 2017 Methods Update Rule | 0.027 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.818 | As described in 2017 Methods Update Rule | 0.066 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 1.5 | ±30% | NA | NA | 1 | 0.5 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes: Allow the use of collision cell technology (KED) | 100000 | No
| Unknown | 2 | ±30% | 2 | NA | 0 | NA | NA | 0.018 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| 80%+ | 5 | ±30% | 0.75 | NA | 1 | 0.7 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | 30% | No | NA | Yes
| NA | 5 | ±30% | 5 | NA | 5 | 0.0213 (Undigested Blank) | As described in method | 0.235 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| 80%+ | 10 | ±30% | 100 | NA | 50 (1-point calibration) | 0.23 | As described in 2017 Methods Update Rule | 0.27 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 10 | ±10% | NA | NA | 1 | NA | NA | 0.0594 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 10 | NA | NA | NA | 10 | 1.8 | As described in 2017 Methods Update Rule | 1.3 | NA | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | NA | 0.1 | 0.002 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work. | None | No | NA | Yes

## Thallium
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to DLR of 0.1 ug/L? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.1 ug/L | 6e. Equipment Needed for DLR of 0.1 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.8 | NA | 0.1 | 85-115% | 0.1 | Yes | 0.1 | 0.03 | As described in 2017 Methods Update Rule | 0.02 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| NA | 0.1 | ±70% | 0.1 | Yes | 0.1 | 0.015 | As described in 2017 Methods Update Rule | 0.016 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| NA | 0.1 | ±30% | 0.11 | Yes | 0.1 | 0.0141 | As described in 2017 Methods Update Rule | 0.0134 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 80%+ | 0.2 | ±25% | 0.2 | No | 0.2 | 0.024 | As described in 2017 Methods Update Rule | 0.053 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 0.2 | ±30% | 0.2 | Yes | 0.1 | 0.018 | As described in 2017 Methods Update Rule | 0.021 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 80%+ | 0.5 | ±30% | 0.05 | Yes | 1 | 0.05 | As described in 2017 Methods Update Rule | 0.2 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | 30% | No | NA | Yes
| 80%+ | 0.5 | ±15% | 0.5 | Yes | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 30% | No | NA | Yes
| NA | 0.5 | ±30% | NA | No | 1 | 0.1 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | greater than 30% | Yes | 100000 | No
| NA | 1 | ±30% | 1 | No | 1 | 0.019 (Undigested Blank) | As described in method | 0.0394 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 1 | ±30% | 1 | Yes | 0 | NA | NA | 0.022 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | Yes | 1 | 0.014 | As described in 2017 Methods Update Rule | 0.014 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | No | 1 | 0.95 | As described in 2017 Methods Update Rule | 0.59 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | No
| NA | 1 | ±30% | 200 | No | 50 (1-point calibration) | 0.046 | As described in 2017 Methods Update Rule | 0.054 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 80%+ | 1 | 74-119% | 1 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 1 | 0.084 | As described in 2017 Methods Update Rule | 0.054 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | NA | NA | No | 1 | 0.48 | As described in 2017 Methods Update Rule | 0.33 | As described in 2017 Methods Update Rule | 200 | Suite | NA | greater than 30% | Yes | NA | No
| NA | 1 | NA | NA | No | 1 | 0.22 | As described in 2017 Methods Update Rule | 0.64 | As described in 2017 Methods Update Rule | 15 | Suite | NA | 10% | No | NA | Yes
| NA | 3 | ±10% | NA | NA | NA | NA | NA | 0.0356 | As described in method | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | PQL = 0.1 ug/L | 70-130% (MS/MSD) | Yes; LCS = 85-115% | Yes | 0.05 | 0.08 | As described in 2017 Methods Update Rule | 0.09 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Vanadium
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to the target MRL? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 0.5 ug/L | 6e. Equipment Needed for DLR of 0.5 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | NA | 100 | 90-110% | 80-120% | NA | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | No
EPA 200.8 | 80%+ | 0.5 | ±15% | 0.5 | NA | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | 10% | No | NA | Yes
| NA | 0.5 | ±30% | 0.5 | NA | 0.1 | 0.052 | As described in 2017 Methods Update Rule | 0.155 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| NA | 0.5 | ±30% | 0.55 | NA | 0.5 | 0.12 | As described in 2017 Methods Update Rule | 0.132 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| NA | 1 | ±70% | 1 | NA | 1 | 0.025 | As described in 2017 Methods Update Rule | 0.015 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| NA | 1 | ±30% | 1 | NA | 1 | 0.0253 (Undigested Blank) | As described in method | 0.071 (Undigested Sample) | As described in method | NA | Suite | NA | greater than 30% | Yes; EPA 200.8 only permits the use of the ICPMS, which we have. If we had to go to this new level we would need EPA approval to use collision cell technology and/or the ICP with triple quad MS. | $400,000 (cost of an ICP with triple quad MS) | Yes
| NA | 1 | ±30% | 1 | NA | 0 | NA | NA | 0.008 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 1 | ±30% | 1 | NA | 1 | 0.021 | As described in 2017 Methods Update Rule | 0.025 | As described in 2017 Methods Update Rule | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.05 | As described in 2017 Methods Update Rule | 0.072 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | 85-115% | Unknown | NA | 1 | 0.7 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 50 | Single | Varies | greater than 30% | Yes | 300000 | Yes
| 80%+ | 2 | ±25% | 2 | NA | 1 | 0.022 | As described in 2017 Methods Update Rule | 0.087 | As described in 2017 Methods Update Rule | 170 | Suite | 15% | 25% | Yes | 140000 | Yes
| NA | 2.5 | ±30% | NA | NA | 1 | 0.5 | As described in 2017 Methods Update Rule | 0.1 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | 25% | No | 25000 | Yes
| NA | 3 | ±30% | 200 | NA | 50 (1-point calibration) | 0.63 | As described in 2017 Methods Update Rule | 0.63 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| 80%+ | 3 | 87-126% | 3 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 3 | 0.31 | As described in 2017 Methods Update Rule | 1.02 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| NA | 3 | ±10% | NA | NA | 1 | NA | NA | 0.556 | As described in 2017 Methods Update Rule | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
| NA | 0.5 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | NA | 0.5 | 2 | As described in 2017 Methods Update Rule | 0.3 | As described in method | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes

## Zinc
1a. Method | 1b. Method Usage | 2a. MRL (ug/L) | 2b. Recovery Range | 2c. MRL (99% Confidence with Recovery btwn 70-130%) (ug/L) | 3a. Can detect to the target MRL? | 3b. Lowest Calibration Point (ug/L) | 4a. MDL(b) (ug/L) | 4b. Determination Procedure for 4a | 5a. MDL(s) (ug/L) | 5b. Determination Procedure for 5a | 6a. Analysis Cost | 6b. Suite or Single Element | 6c. Pricing Discount | 6d. Percent Cost Increase Expected from DLR of 50 ug/L | 6e. Equipment Needed for DLR of 50 ug/L | 6f. Equipment Cost | 6g. Maintain accreditation if DLR was lowered?
EPA 200.7 | Unknown | 10 | ±30% | 10 | NA | 0 | 0.75 | 2016 EPA MDL Procedure | 0.74 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 10 | ±30% | 10 | NA | 10 | 2.017 | As described in 2017 Methods Update Rule | 1.254 | As described in method | $35 (Estimated for single metal) | Suite | NA-Lab supports tests from it's primary organization's Water Treatment and other related processes. No external clients. | less than 5% | No | NA | Yes
| 20% | 10 | 85-115% | Unknown | NA | 10 | 10 | As described in 2017 Methods Update Rule | 10 | As described in 2017 Methods Update Rule | 40 | Single | Varies | NA | No | NA | Yes
| 80%+ | 20 | ±25% | 2 | NA | 2-point calibration: 0 and 2,000 | 1.2 | As described in 2017 Methods Update Rule | 8 | As described in 2017 Methods Update Rule | $175 (includes Anions, TDS, EC) | Suite | 15% | 25% | Yes | 140000 | Yes
| 80%+ | 30 | ±30% | NA | NA | 10 | 10 | As described in 2017 Methods Update Rule | 7 | As described in 2017 Methods Update Rule | 25 | Suite | Yes | 25% | No | 25000 | Yes
| 0.2 | 50 | ±30% | 50 | NA | 5,000 (1-point calibration) | 1.02 | As described in 2017 Methods Update Rule | 0.74 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 75% | 50 | ±30% | 55 | NA | 50 | 1.24 | As described in 2017 Methods Update Rule | 1.2 | As described in 2017 Methods Update Rule | $155.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 50% | 50 | ±30% | 200 | NA | 25,000 (1-point calibration) | 3.4 | As described in 2017 Methods Update Rule | 4.8 | As described in 2017 Methods Update Rule | $25 (Discounted for suite of metals - see table) | Single | $15 - $22.50 (Discounts based on client revenue - see table) | 30% | Yes | 150000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 50 | 90-110% | 80-90% | Yes | 40 | 40 | As described in 2017 Methods Update Rule | NA | NA | 20 | Suite | Yes | less than 5% | Yes | 150000 | No
| NA | 50 | ±15% | NA | NA | 5000 | NA | NA | 9.86 | As described in 2017 Methods Update Rule | 20 | Single | It depends on proj scope | greater than 30% | NA | NA | Yes
EPA 200.8 | 80%+ | 0.5 | ±15% | 0.5 | NA | 0.5 | NA | NA | 0.1 | As described in 2017 Methods Update Rule | 12 | Suite | NA | less than 5% | No | NA | Yes
| NA | 1 | ±50% | 10 | Yes | 1 | 0.236 | As described in 2017 Methods Update Rule | 0.32 | As described in 2017 Methods Update Rule | Unknown | Suite | NA | less than 5% | No | NA | Yes
| 30% | 3 | ±30% | 3.3 | NA | 3 | 1.27 | As described in 2017 Methods Update Rule | 0.316 | As described in 2017 Methods Update Rule | $254.10 (minimum) | Suite (Customer may request single analyte within the suite) | None | less than 5% | No | NA | Yes
| 80%+ | 5 | ±30% | 0.7 | NA | 1 | 3 | As described in 2017 Methods Update Rule | 0.7 | As described in 2017 Methods Update Rule | 32 | Single | Yes, 10% discount | less than 5% | No | NA | Yes
| Unknown | 5 | ±30% | 5 | NA | 0 | NA | NA | 0.45 | 2016 EPA MDL Procedure | $50 for 1 analayte, $60 for 2, +$12 for each additional | Whatever elements are selected | It depends | 0% | No | NA | Yes
| NA | 5 | NA | NA | NA | 5 | 1.04 | As described in 2017 Methods Update Rule | 1.4 | As described in 2017 Methods Update Rule | 200 | Suite | NA | less than 5% | No | NA | No
| 75% | 5 | 85-115% | Unknown | NA | 5 | 1 | As described in 2017 Methods Update Rule | 1 | As described in 2017 Methods Update Rule | 50 | Single | Varies | NA | No | NA | Yes
| NA | 10 | ±70% | 10 | NA | 10 | 0.103 | As described in 2017 Methods Update Rule | 0.724 | As described in 2017 Methods Update Rule | 20 | Single | NA | No Idea | Likely | No Idea | Maybe
| 80%+ | 10 | ±30% | 10 | NA | 0.5 | 0.704 | As described in 2017 Methods Update Rule | 0.804 | As described in 2017 Methods Update Rule | 21 | Suite | Yes (varies) | 5% | No | NA | Yes
| 20% | 10 | ±30% | NA | NA | 1 | NA | As described in 2017 Methods Update Rule | NA | As described in 2017 Methods Update Rule | 25 | Suite | Yes | 25% | No | 25000 | Yes
| NA | 20 | ±30% | 20 | NA | 20 | 0.0664 (Undigested Blank) | As described in method | 0.513 (Undigested Sample) | As described in method | NA | Suite | NA | NA | No | NA | Yes
| 80%+ | 20 | 33-167% | 20 (MRL is 50-150%, do not have statistics at 70-130%) | Yes | 20 | 2.24 | As described in 2017 Methods Update Rule | 1.28 | As described in 2017 Methods Update Rule | 15 | Suite | NA | less than 5% | No | NA | Yes
| 50% | 50 | ±30% | 200 | NA | 50 (1-point calibration) | 0.22 | As described in 2017 Methods Update Rule | 1.9 | As described in 2017 Methods Update Rule | $30 (Discounted for suite of metals - see table) | Single | $25 - $28 (Discounts based on client revenue - see table) | 30% | Yes | 250000 | Yes (Assuming clients are agreeable to higher cost of analysis)
| NA | 50 | NA | NA | NA | 50 | 31 | As described in 2017 Methods Update Rule | 22 | As described in 2017 Methods Update Rule | 15 | Suite | NA | 10% | No | NA | Yes
| NA | 5 (PQL) | 70-130% (MS/MSD) | Yes; LCS = 85-115% | NA | 1 | 0.8 | As described in 2017 Methods Update Rule | 2 | As described in 2017 Methods Update Rule | $40/metal | Both single and suite | Yes, depends on scope of work | None | No | NA | Yes