## TOC
SWRCB. (2022c). Laboratory Survey Responses: December 2022. State Water Resources Control Board.
Table of Contents
Antimony
Arsenic
Beryllium
Cadmium
Lead
Manganese
Mercury
Thallium
Iron
Miscellaneous

## Antimony
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.8 | 1 | No change | 0.1 | 0.033786 | 0.046357 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | 15% | NA | No | 150 | NA
|  | No change | 0.1 | 0.04 | 0.04 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.5 | 0.005 | 0.4 | EPA 821-R-16-006, December 2016 | No | 2 | No | NA | NA | We would not pursue ELAP accreditation for this contaminant | NA | NA | NA | NA | NA | Yes | 3
|  | No change | 0.5 | 0.005 | 0.4 | EPA 821-R-16-006, December 2016 | No | 2 | No | NA | NA | We would not pursue ELAP accreditation for this contaminant | NA | NA | NA | NA | NA | Yes | 3
|  | NA | 0.5 | 0.02 | 0.15 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | MDL study at lower spiking concentration, SOP/LIMS update, possibly method development | 6 | NA | NA | NA | NA | No | 0 for drinking water | NA
|  | NA | 0.5 | 0.098 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | We are already at a lower RL for NPDES monitoring | No | 20 | NA
|  | NA | 0.5 | NA | 0.036 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | NA | NA
|  | No change | 0.5 | NA | 0.57 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | Costs depend on what the markey is doing at any given point in time. Carrier gases continue to fluctuate as do acid costs. Labor costs have a diredt impact on analytical costs. | No | 50 | NA
|  | No change | 1 | 0.0145 | 0.0559 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.02 | 0.167 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 | NA | NA | NA | 8 | NA
|  | No change | 1 | 0.06 | 0.024 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | NA | NA
|  | No change | 1 | 0.12043 | 0.09482 | As described in 2017 Methods Update Rule | No | Typically 5 times the low limit is considered reasonable to meet 30% regularly | No | NA | NA | NA | NA | unknown, this financial information comes from high corporate levels |  | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.12043 | 0.09482 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.2 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Current reporting limit equal to proposed DLR | NA | No impact | No impact | No impact | NA | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | 1 | 0.3861 | 0.461 | As described in 2017 Methods Update Rule | No | 2 | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 30 | NA
|  | No change | 1 | 0.5 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | Analyst time and instrument maintenance would increase; 20% cost increase |  | 20 to 50% | NA | No | 125 | NA
|  | No change | 2 | 0.022 | 0.027 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | No | 11 | NA
|  | NA | 6 | 0.0412 | 0.342 | 40CFR Part 136, App B | No | 6 | No | additional study | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 1 | 0.06 | 0.093 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% | 5% | NA | No | 21 | NA
|  | No change | NA | 0.02 | 0.06 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 61 | NA
|  | None, subcontract | NA | NA | NA | NA | Yes | NA | No | NA | NA | Method development | 6 | 5% | 5% | NA | NA | No | 5 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
EPA 200.9 | 1 | EPA 200.8 | NA | NA | 2 | As described in method | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Yes | 3 | NA

## Arsenic
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.8 | 0.5 | No change | 0.1 | 0.0411 | 0.11191 | As described in 2017 Methods Update Rule | No | +/- 50% unless reaction gas is permitted | No | NA | NA | Permission to use reaction gas ICPMS for drinking water samples | NA | NA | NA | 15% | NA | No | 340 | NA
|  | No change | 0.2 | 0.09 | 0.09 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 | NA | NA | NA | 8 | NA
|  | No change | 0.5 | 0.08 | 0.1 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 100 | NA
|  | No change | 0.5 | 0.08 | 0.1 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 100 | NA
|  | No change | 0.5 | 0.3 | 0.2 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 0 | NA | No | 400 to 600 | NA
|  | NA | 0.5 | 0.37 | 0.35 | As described in 2017 Methods Update Rule | No | NA | No | ICPMS using Collision Cell | NA | Regulation Update to use Collision cell for drinking water analysis; Method Development; analyst training | 6 | NA | NA | NA | NA | No | 20 | NA
|  | NA | 0.5 | 0.7 | 0.53 | As described in 2017 Methods Update Rule | No | NA | No | We already have collision cell technology, but EPA's 2006 memo prohibits this for drinking water compliance (https://www.waterboards.ca.gov/drinking_water/certlic/labs/documents/ELAPCollisionCellDW200_8.pdf) | NA | EPA regulatory change to allow collision/reaction cells for EPA 200.8 analysis of drinking water | NA | NA | NA | NA | NA | Yes | 0 for drinking water | Chloride-argon interference causes false positive sample results unless a collision/reaction cell is used for EPA 200.8. My understanding is that collision/reaction cells are prohibited by EPA for drinking water compliance testing.
|  | NA | 0.5 | NA | 0.095 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 30 | NA
|  | No change | 0.5 | NA | 0.27 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 50 | NA
|  | No change | 1 | 0.0875 | 0.1636 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | Analyst time and instrument maintenance would increase; 20% cost increase |  | 20 to 50% | NA | No | 125 | NA
|  | No change | 1 | 0.128 | 0.087 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | No | 60 | NA
|  | No change | 1 | 0.188 | 0.385 | As described in 2017 Methods Update Rule | No | 1 | Yes | Regulatory approval of Collision Cell technology built within the current instrumentation. Or purchase of new instrumentation. | If use of collison cell technology is allowed, no additional cost, if new instrument purchase required then approx $100K | Utilize collision cell technology; change calibration concentrations; update SOP and Standard Prep form; perform MDL studies; train staff. | 6 to 12 | If new equipment purchase is needed, then 30% increase in capital expenditure in the year when the new instrument is purchased. | Only the instrument purchase cost increase. Operating cost expected to remain same. | NA: Utility lab, internal customers only | NA | Yes | 30 | If collision cell is acceptable, then 0.5 ppb may be possible. However, 1 ppb is more realistic.
|  | No change | 1 | 0.24 | 0.16 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development | 3 to 6 | 2% | No impact | 2% | Primary factors include staff testing existing instrument capabilities, potentially investing in new custom calibration standards, etc. to meet proposed DLR | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | 1 | 0.4103 | 0.49123 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.4103 | 0.49123 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.4746 | 0.8174 | As described in 2017 Methods Update Rule | No | 2 | Yes | ICP-MS with DRC | 250000 | staff training, method development | 6 | 200% | NA | NA | NA | Yes | 475 | DRC is currently not acceptable for drinking water accreditation.
|  | NA | 2 | 0.814 | 0.797 | 40CFR Part 136, App B | No | 2 | No | additional study | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 1 | 0.087 | 0.13 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% | 5% | NA | No | 142 | NA
|  | No change | NA | 0.3 | 0.3 | As described in 2017 Methods Update Rule | No | ~1 | Yes | Equipment that can achieve the DLR. | Unknown. If new instrument is needed, ~$350,000. | Method development and potentially staff training. | 4 to 6 | 30 to 35% if new instrument | 10 to 20% | 20% | NA | NA | NA | All best estimates. Please consider collison cell as an alternative.
|  | No change | NA | NA | NA | NA | No | NA, 56% of the time currently | No | NA | NA | Investigation into cause | 6 | NA | NA | NA | NA | No | 150 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 388 | NA
|  | None, subcontract | NA | NA | NA | NA | No | 5 | Yes | He plasma, ICPMS. | 30000 | New plasma, equipment, training, development, infrastructure | 24 | 200% | 100% | NA | NA | Yes | 10 | Can't be done without considerable investment
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
EPA 200.9 | 0.5 | EPA 200.8 | NA | NA | 0.3 | As described in method | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Yes | 20 | NA
SM 3113B | 0.5 | EPA 200.8 | 2 | 0.5 | 0.6 | As described in 2017 Methods Update Rule | No | 2 | Yes | ICP-MS | 400000 | Research instrument choices including attending vendor presentations, prepare request for proposals (RFPs) with instrument specifications, evaluate RFPs received, select vendor, prepare Board agenda item, obtain Board approval to purchase instrument and any needed appertunances (compressed gases, ventilation equipment, power requirements), order/ purchase instrument and appurtenances, install appertunances and instrument, vendor to perform initial instrument startup, method development, demonstration of capability, vendor to train laboratory staff, lab to perform proficiency testing, request and receive quotes for third party assessment agency (TPA) to perform laboratory onsite assessment, prepare corrective action plan (CAP) as needed as result of onsite assessment, submit application to ELAP for accreditation amendment along with appplication fee and onsite assessment CAP, receive amended accreditation from ELAP | 36 | 100% | 4.2% | 65% | NA | No | 63 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continued Staff training - optimizing instrument | 1 | 0 | 0 | 0 | NA | No | 6 samples quarterly | NA

## Beryllium
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.7 | 0.5 | EPA 200.8 | 10 | NA | 0.1 | As described in method | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Yes | 3 | NA
|  | EPA 200.8 | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Yes | 0 | NA
|  | EPA 200.8 | NA | NA | NA | NA | No | 1 | No | NA | NA | Method development | 12 | 10% | 10% | NA | NA | No | 1 | NA
EPA 200.8 | 0.5 | No change | 0.1 | 0.00426 | 0.017894 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | 15% | NA | No | 140 | NA
|  | NA | 0.1 | 0.007 | 0.051 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | MDL study at lower spiking concentration, SOP/LIMS update, possibly method development | 6 | NA | NA | NA | NA | No | 0 for drinking water | NA
|  | No change | 0.1 | 0.02 | 0.08 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.2 | 0.07 | 0.0299 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 | NA | NA | NA | 8 | NA
|  | NA | 0.5 | 0.013 | 0.051 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 20 | NA
|  | No change | 0.5 | 0.03 | 0.1 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 3 | NA
|  | No change | 0.5 | 0.03 | 0.1 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 3 | NA
|  | NA | 0.5 | NA | 0.064 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | NA | NA
|  | No change | 0.5 | NA | 0.15 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 50 | NA
|  | No change | 1 | 0.002 | 0.022 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | No | 11 | NA
|  | NA | 1 | 0.0067 | 0.0635 | 40CFR Part 136, App B | No | 1 | No | additional study | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 1 | 0.008 | 0.012 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 30 | NA
|  | No change | 1 | 0.0195 | 0.0552 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.0768 | 0.2786 | As described in 2017 Methods Update Rule | No | 2 | Yes | ICP-MS with DRC | 250000 | staff training, method development | 6 | 200% | NA | NA | NA | Yes | 60 | DRC is currently not acceptable for drinking water accreditation.
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | Analyst time and instrument maintenance would increase; 20% cost increase |  | 20 to 50% | NA | No | 125 | NA
|  | No change | 1 | 0.13 | 0.11 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development | 3 to 6 | 2% | No impact | 2% | Primary factors include staff testing existing instrument capabilities, potentially investing in new custom calibration standards, etc. to meet proposed DLR | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | 1 | 0.17702 | 0.12008 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.17702 | 0.12008 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.04 | 0.073 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% | 5% | NA | No | 20 | NA
|  | No change | NA | 0.005 | 0.04 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 60 | NA
|  | None, subcontract | NA | NA | NA | NA | No | 1 | No | NA | NA | Method development | 12 | 10% | 10% | NA | NA | No | 1 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SM 3113B | 0.5 | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continue Staff training - optimizing instrument | 1 | 0 | 0 | 0 | NA | No | 6 per quarter | NA

## Cadmium
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.7 | 0.5 | EPA 200.8 | 20 | NA | 0.2 | As described in method | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Yes | 2 | NA
EPA 200.8 | 0.5 | No change | 0.1 | 0.004 | 0.2 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.1 | 0.0257 | 0.11161 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | 15% | NA | No | 300 | NA
|  | No change | 0.2 | 0.07 | 0.257 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 | NA | NA | NA | 8 | NA
|  | NA | 0.25 | 0.015 | 0.058 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 20 | NA
|  | No change | 0.5 | 0.014 | 0.025 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | No | 11 | NA
|  | NA | 0.5 | 0.015 | 0.076 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | MDL study at lower spiking concentration, SOP/LIMS update, possibly method development | 6 | NA | NA | NA | NA | No | 0 for drinking water | NA
|  | No change | 0.5 | 0.024639 | 0.069853 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 0.5 | 0.024639 | 0.069853 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 0.5 | 0.1 | 0.3 | EPA 821-R-16-006, December 2016 | No | 1.5 | No | NA | NA | We would not pursue ELAP accreditation for this contaminant | NA | NA | NA | NA | NA | Yes | 3 | NA
|  | No change | 0.5 | 0.1 | 0.3 | EPA 821-R-16-006, December 2016 | No | 1.5 | No | NA | NA | We would not pursue ELAP accreditation for this contaminant | NA | NA | NA | NA | NA | Yes | 3 | NA
|  | NA | 0.5 | NA | 0.083 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | NA | NA
|  | No change | 0.5 | NA | 0.17 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 50 | NA
|  | No change | 1 | 0.004 | 0.026 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 30 | NA
|  | No change | 1 | 0.01 | 0.076 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% | 5% | NA | No | 24 | NA
|  | No change | 1 | 0.0133 | 0.145 | As described in 2017 Methods Update Rule | No | 2 | Yes | ICP-MS with DRC | 250000 | staff training, method development | 6 | 200% | NA | NA | NA | Yes | 45 | DRC is currently not acceptable for drinking water accreditation.
|  | NA | 1 | 0.0196 | 0.0686 | 40CFR Part 136, App B | No | 1 | No | additional study | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 1 | 0.0218 | 0.0554 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | Analyst time and instrument maintenance would increase; 20% cost increase |  | 20 to 50% | NA | No | 125 | NA
|  | No change | 1 | 0.27 | 0.23 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development | 3 to 6 | 2% | No impact | 2% | Primary factors include staff testing existing instrument capabilities, potentially investing in new custom calibration standards, etc. to meet proposed DLR | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 66 | NA
|  | None, subcontract | NA | NA | NA | NA | No | 1 | No | NA | NA | Method development. | 12 | 10% | 10% | NA | NA | No | 5 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SM 3113B | 0.5 | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continued Staff training - optimizing instrument | 1 | 0 | 0 | 0 | NA | No | 6 per quarter | NA

## Iron
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Cost of analysis using method based on current regulation ($) | Is the sample analysis price based on a single element or suite of elements? | If available, pricing discount for contracted or regular clients (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.7 | 100 | EPA 200.5 | 0 | 20.1 | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | 30 | Yes | NA | NA | NA | NA | 2 | NA
|  | No change | 5 | NA | 3 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Current reporting limit below proposed DLR | NA | No impact | No impact | Unknown | Suite | NA | 0 | NA | No | 60 | NA
|  | No change | 10 | 0.562 | 0.7 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA, Utility lab | NA, Utility lab | NA | None | NA | No | 30 | NA
|  | No change | 10 | 1.7409 | 1.3989 | As described in 2017 Methods Update Rule | Yes | For No answers: Typically 5 times the low limit is considered reasonable to meet 30% regularly | No | A new instrument. Metals are run as a single suite. If one metal is a problem, a new instrument will need to be purchased, and as evidenced by current data, our instrumentation is already on the borderline with random background in blanks. This may require the entire method be run with HPLC or other ultrapure water, further increasing costs. | NA | None, the existing noted expectation from Column J header is unrealistic. Standard Recovery acceptance criteria is +/- 50% at low levels. 30% is randomly achievable but unreliable on the whole. Truly achieving 30% for any one analyte will simply require new instrumentation, pushing all estimates very high. See most recent EPA metals analysis under UCMR3 (200.8), 4 (200.8), and 5 (200.7), as well as industry practice under all new methodolgy when dealing with low level QC. Low level recovery acceptance criteria is set at +/- 50% as a standard for all current EPA methods both new and updated. | Several Years as stated, this is a large CapEx request | unknown, this financial information comes from high corporate levels |  | NA | NA | NA | NA | NA | No | >1000 | If the state follows current EPA trends and expects 50-150 recovery criteria at DLR (which tends to be at or near our low calibration point) then these limits are relatively easy to achieve. At 30%, this is a huge cost sunk into metals analysis to achieve what are special, California only limits, be it on equipment cost or high purity reagent cost (with the reagent cost, long term, in danger of eclipsing instrument costs). It is strongly recommended the state look at the laboratory approval manuals for UCMR 3/4/5 related to metals analysis, and the QC limits contained therein.
|  | No change | 10 | 1.7409 | 1.3989 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | NA | 10 | 2.5 | 3.9 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | We are already at a lower RL for NPDES monitoring | No | 20 | NA
|  | No change | 10 | 10 | 30 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | 0 | Analyst time and instrument maintenance would increase; 20% cost increase |  | 26 | Single | Yes, up to 10% discount; $23.40 | 20% increase; $31.20 | NA | No | NA | <-- 175
|  | No change | 30 | 7 | 11.8 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | 25 | Single | NA | 25% | Discount based on total revenue | No | 200 | "Lowest calibration point" is the CRL verified daily (+/- 40%) of a single-point calibration curve.
|  | NA | 50 | NA | 20 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | Single | NA | 0 | NA | No | 5 | NA
|  | NA | 100 | 23.4 | 23.6 | 40CFR Part 136, App B | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 100 | NA | 6.83 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | 100 | NA | 4.8 | As described in method | Yes | NA | Yes | NA | $30,000 to $60,000 | Staff training, method development | 24 | 100% | 25% | 30 | Single | 10% | 5% | NA | No | 40 | Filled out what I could with the time I could allocate to the survey.
|  | No change | 10000 | 6.297 | 5.725 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | NA | NA | Yes | 0 | Pricing depends greatly on other work. Sometimes metals analysis are used as loss leaders. | No | 310 | NA
|  | No change | NA | 1.6 | 0.9 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | 35 | Single | NA | NA | NA | No | NA | NA
|  | No change | NA | 1.91 | 2.34 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | NA | NA | NA | No | 150 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 100 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 403 | NA
|  | NA | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | NA | NA | NA | N | 5 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
EPA 200.8 | 100 | No change | 0.5 | 0.06 | 0.6 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | 26.79 | Single | NA | NA | NA | NA | 3 | NA
|  | No change | 0.5 | 0.06 | 0.6 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | 26.79 | Single | NA | NA | NA | NA | 3 | NA
|  | No change | 1 | 5 | 10 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | ELAP would need to allow Fe by EPA200.8 for drinking water | 0 | Analyst time and instrument maintenance would increase; 20% cost increase |  | 26 | Single | Yes, up to 10% discount; $23.40 | 20% increase; $31.20 | NA | No | 26- no regulatory | NA
|  | NA | 5 | 4.6 | 4.9 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | SOP/LIMS update | 1 | NA | NA | NA | NA | No | 0 | NA | No | 0 for drinking water | NA
|  | No change | 25 | 4 | 8 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 29 | Suite | Yes | 0 | Discounting depends on scope; single element cost = $52 | No | 400 to 600 | NA
|  | NA | 50 | NA | 3.632 | As described in method | Yes | NA | NA | NA | NA | NA | NA | NA | NA | 30 | Single | 10% | NA | wastewater samples | No | 95 | NA
|  | No change | 100 | NA | NA | ELAP doesn't offer Accredited fields of testing to Iron by EPA method 200.8. Therefore, lab uses the lowest calibration level as detection level. The iron results were for research purpose only. | Yes | NA | No | ICP-MS | NA | ELAP shall offer accreditation to EPA200.8 Iron for drinking water test in California | NA | 0 | 0 | 7 | Suite of elements | NA | 0 | NA | No | 30 | NA
SM 3111B | 100 | No change | 100 | 6.43 | 20.4 | as described in 2017 Methods Update Rule | Yes | NA | No | None | NA | None | 0 | 0 | 0 | NA | NA | NA | NA | NA | No | 8 | NA
|  | EPA 200.7 | 500 | 46 | 26 | As described in 2017 Methods Update Rule | Yes | na | No | na | na | na | na | na | na | NA | NA | NA | NA | NA | No | 6 to 10 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continued Staff training - optimizing instrument | 1 | 0 | 0 | $15 (what we charge) | Single | NA | 0 | NA | No | 6 per quarter | NA
NA | 100 | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 30 | Yes | NA | NA | NA | NA | 2 | NA

## Lead
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Cost of analysis using method based on current regulation ($) | Is the sample analysis price based on a single element or suite of elements? | If available, pricing discount for contracted or regular clients (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.5 | 1 | No change | 0 | 0.92 | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | 30 | Yes | NA | NA | NA | NA | 2 | NA
EPA 200.8 | 0.5 | No change | 0.1 | 0.011 | 0.009 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | No impact | No impact |  |  |  | NA | NA | No | 30 samples every 3 years | NA
|  | No change | 0.1 | 0.024866 | 0.030981 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA |  |  |  | 15% | NA | No | 530 | NA
|  | No change | 0.1 | 0.03 | 0.04 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 |  |  |  | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.2 | 0.05 | 0.05 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 |  |  |  | NA | NA | NA | 8 | NA
|  | No change | 0.5 | 0.003 | 0.06 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 3 | NA
|  | No change | 0.5 | 0.003 | 0.06 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 3 | NA
|  | NA | 0.5 | 0.0069 | 0.0451 | 40CFR Part 136, App B | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
|  | No change | 0.5 | 0.012 | 0.011 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA |  |  |  | NA | NA | No | 170 | NA
|  | NA | 0.5 | 0.018 | 0.215 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 20 | NA
|  | NA | 0.5 | 0.04 | 0.11 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | MDL study at lower spiking concentration, SOP/LIMS update, possibly method development | 6 | NA | NA |  |  |  | NA | NA | No | 0 for drinking water | NA
|  | No change | 0.5 | 0.29164 | 0.081525 | As described in 2017 Methods Update Rule | [12 empty] | NA | NA | No | >1000 | NA | bar: Yes to NA
|  | No change | 0.5 | 0.29164 | 0.081525 | As described in 2017 Methods Update Rule | [12 empty] | NA | NA | No | >1000 | NA | bar: Yes to NA
|  | NA | 0.5 | NA | 0.041 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 |  |  |  | NA | NA | No | 5 | NA
|  | No change | 0.5 | NA | 0.23 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 |  |  |  | NA | NA | No | 50 | NA
|  | No change | 1 | 0.01 | 0.031 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 30 | NA
|  | No change | 1 | 0.0208 | 0.0387 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 |  |  |  | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.0309 | 0.1074 | As described in 2017 Methods Update Rule | No | 2 | Yes | ICP-MS with DRC | 250000 | staff training, method development | 6 | 200% | NA |  |  |  | NA | NA | Yes | 350 | DRC is currently not acceptable for drinking water accreditation.
|  | No change | 1 | 0.07 | 0.083 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% |  |  |  | 5% | NA | No | 98 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | Analyst time and instrument maintenance would increase; 20% cost increase |  |  |  |  | 20 to 50% | NA | No | 125 | NA
|  | No change | 1 | 0.36 | 0.24 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development | 3 to 6 | 2% | No impact |  |  |  | 2% | Primary factors include staff testing existing instrument capabilities, potentially investing in new custom calibration standards, etc. to meet proposed DLR | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | NA | 0.03 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 400 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 582 | NA
|  | None, subcontract | NA | NA | NA | NA | Yes | NA | No | NA | NA | Method development. | 6 | 5% | 5% |  |  |  | NA | NA | No | 50 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
| 1 | No change | 0.1 | 0.011 | 0.009 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | No impact | No impact | NA | NA | NA | NA | NA | No | 30 samples every 3 years | NA
|  | No change | 0.1 | 0.024866 | 0.030981 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | NA | NA | Yes | 0 | Pricing depends greatly on other work. Sometimes metals analysis are used as loss leaders. | No | 530 | NA
|  | No change | 0.1 | 0.03 | 0.04 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 29 | Suite | Yes | 0 | Discounting depends on scope; single element cost = $52 | No | 400 to 600 | NA
|  | No change | 0.2 | 0.05 | 0.0307 | as described in 2017 Methods Update Rule | Yes | NA | No | None | NA | None | 0 | 0 | 0 | NA | NA | NA | NA | NA | No | 8 | NA
|  | No change | 0.5 | 0.003 | 0.06 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | 26.79 | Single | NA | NA | NA | NA | 3 | NA
|  | No change | 0.5 | 0.003 | 0.06 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | 26.79 | Single | NA | NA | NA | NA | 3 | NA
|  | NA | 0.5 | 0.0069 | 0.0451 | 40CFR Part 136, App B | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 0.5 | 0.012 | 0.011 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | NA | NA | NA | No | 170 | NA
|  | NA | 0.5 | 0.018 | 0.21 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 20 | NA
|  | NA | 0.5 | 0.04 | 0.11 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | MDL study at lower spiking concentration, SOP/LIMS update | 3 | NA | NA | NA | NA | No | 0 | NA | No | 0 for drinking water | NA
|  | No change | 0.5 | 0.29164 | 0.081525 | As described in 2017 Methods Update Rule | No | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 0.5 | 0.29164 | 0.081525 | As described in 2017 Methods Update Rule | No | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | NA | 0.5 | NA | 0.041 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | Single | NA | 0 | NA | No | 5 | NA
|  | No change | 0.5 | NA | 0.23 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | 1 | 0.01 | 0.031 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA, Utility lab | NA, Utility lab | NA | None | NA | No | 30 | NA
|  | No change | 1 | 0.0208 | 0.0387 | As described in 2017 Methods Update Rule | Yes | NA | No | ICP-MS | NA | Staff training | 3 | 0 | 0 | 7 | Suite | NA | 0 | NA | No | 4 | NA
|  | No change | 1 | 0.0309 | 0.1074 | As described in 2017 Methods Update Rule | No | 2 | No | NA | NA | NA | NA | NA | NA | 35 | Single | NA | 25% | Discount based on total revenue | No | 350 | "Lowest calibration point" is the CRL verified daily (+/- 40%) of a single-point calibration curve.
|  | No change | 1 | 0.07 | 0.083 | As described in method | Yes | NA | No | NA | NA | NA | NA | NA | NA | 30 | Single | 10% | NA | NA | No | 96 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | 0 | Analyst time and instrument maintenance would increase; 20% cost increase |  | 26 | Single | Yes, up to 10% discount; $23.40 | 20% increase; $31.20 | NA | No | 175 | NA
|  | No change | 1 | NA | 0.36 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Current reporting limit equal to proposed DLR | NA | No impact | No impact | Unknown | Suite | NA | 0 | NA | No | 60 | NA
|  | No change | NA | 0.03 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | 35 | Single | NA | NA | NA | No | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 400 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 582 | NA
|  | NA | NA | NA | NA | NA | Yes | NA | No | NA | NA | Method development. | 6 | 5% | 5% | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
EPA 200.9 | 0.5 | EPA 200.8 | NA | NA | 0.9 | As described in method | [12 empty] | NA | NA | Yes | 4 | NA | bar: Yes to NA
| 1 | EPA 200.8 | 5 | NA | 0.9 | As described in method | No | 5 | No | NA | NA | NA | NA | NA | 10% | 30 | Single | 10% | NA | NA | Yes | 6 | NA
SM 3111B | 0.5 | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continued Staff training - optimizing instrument | 1 | 0 | 0 |  |  |  | 0 | NA | No | 6 per quarter | NA
|  | EPA 200.8 | 5 | 0.8 | 0.6 | As described in 2017 Methods Update Rule | No | 1.5 | Yes | ICP-MS | 400000 | Research instrument choices including attending vendor presentations, prepare request for proposals (RFPs) with instrument specifications, evaluate RFPs received, select vendor, prepare Board agenda item, obtain Board approval to purchase instrument and any needed appertunances (compressed gases, ventilation equipment, power requirements), order/ purchase instrument and appurtenances, install appertunances and instrument, vendor to perform initial instrument startup, method development, demonstration of capability, vendor to train laboratory staff, lab to perform proficiency testing, request and receive quotes for third party assessment agency (TPA) to perform laboratory onsite assessment, prepare corrective action plan (CAP) as needed as result of onsite assessment, submit application to ELAP for accreditation amendment along with appplication fee and onsite assessment CAP, receive amended accreditation from ELAP | 36 | 100% | 4.2% |  |  |  | 65% | NA | No | 21 | NA
| 1 | EPA 200.8 | 5 | 0.8 | 0.6 | As described in 2017 Methods Update Rule | No | 1.5 | Yes | ICP-MS | 400000 |  | 36 | 100% | 4.2% | 40 | Single | NA | 520% | NA | No | 21 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continued Staff training - optimizing instrument | 1 | 0 | 0 | 15 | Single | NA | 0 | NA | No | 6 per quarter | NA

## Manganese
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Cost of analysis using method based on current regulation ($) | Is the sample analysis price based on a single element or suite of elements? | If available, pricing discount for contracted or regular clients (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.5 | 20 | No change | 0 | 4.138 | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | 30 | Yes | NA | NA | NA | NA | 2 | NA
EPA 200.7 | 10 | No change | 10 | 0.121 | 0.526 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
|  | NA | 10 | 1 | 2 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | 20 | NA
|  | No change | 10 | 2.2 | 1.4 | As described in 2017 Methods Update Rule | No | 20 | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 265 | NA
|  | No change | 10 | 10 | 30 | As described in 2017 Methods Update Rule | No | 30 | Yes | ICPMS | 100000 | Use only 200.8 current method to run Mn, will force samples to be run on 2 methods to fulfill new DLR for all samples | 0 | 5% | 5% |  |  |  | 5% | NA | Yes | 200 | New DLR would not force us to lose certification, rather we would not reapply since the 200.8 method would be the only way to report to the new DLR
|  | NA | 10 | NA | 4 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 |  |  |  | NA | NA | No | NA | NA
|  | EPA 200.8 | 20 | NA | 4.5 | As described in method | No | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | Yes | 13 | NA
|  | EPA 200.8 | 2000 | 0.69 | 0.388 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA |  |  |  | 15% | NA | No | 10 | NA
|  | No change | NA | 0.3 | 0.3 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
|  | EPA 200.8 | NA | NA | NA | NA | [12 empty] | NA | NA | Yes | 50 | NA | bar: Yes to NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 442 | NA
|  | None, subcontract | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | 0 | 0 |  |  |  | NA | NA | No | 10 | NA
| 20 | EPA 200.5 | 0 | 1.3 | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | 30 | Yes | NA | NA | NA | NA | 2 | NA
|  | No change | 10 | 0.121 | 0.526 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA, Utility lab | NA, Utility lab | NA | None | NA | No | 40 | NA
|  | NA | 10 | 1 | 2 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 20 | NA
|  | No change | 10 | 2.2 | 1.4 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | 25 | Single | NA | 25% | Discount based on total revenue | No | 275 | "Lowest calibration point" is the CRL verified daily (+/- 40%) of a single-point calibration curve.
|  | No change | 10 | 5 | 15 | As described in 2017 Methods Update Rule | NA | NA | No | NA | NA | NA | 0 | Analyst time and instrument maintenance would increase; 20% cost increase |  | 26 | Single | Yes, up to 10% discount; $23.40 | 20% increase; $31.20 | NA | No | 175 | NA
|  | NA | 10 | NA | 4 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | Single | NA | 0 | NA | No | 5 | NA
|  | EPA 200.8 | 20 | NA | 4.5 | As described in method | Yes | NA | No | Switch to 200.8 | NA | NA | NA | NA | 10% | 30 | Single | 10% | NA | NA | Yes | 13 | NA
|  | No change | 2000 | 0.69 | 0.388 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | NA | NA | Yes | 0 | Pricing depends greatly on other work. Sometimes metals analysis are used as loss leaders. | No | 10 | NA
|  | No change | NA | 0.3 | 0.3 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | 35 | Single | NA | NA | NA | No | NA | NA
|  | No change | NA | NA | NA | NA | [9 empty] | NA | NA | NA | NA | NA | No | 50 | NA | bar: Yes to NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 622 | NA
|  | NA | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | NA | NA | NA | No | 10 | NA
EPA 200.8 | 10 | No change | 0.1 | 0.06 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 |  |  |  | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.1 | 0.074012 | 0.11569 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA |  |  |  | 15% | NA | No | 230 | NA
|  | No change | 0.2 | 0.06 | 0.0217 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 |  |  |  | NA | NA | NA | 8 | NA
|  | No change | 0.5 | 0.016 | 0.2 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 20 | NA
|  | No change | 0.5 | 0.016 | 0.2 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 20 | NA
|  | NA | 0.5 | 0.03 | 0.12 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | SOP/LIMS update | 6 | NA | NA |  |  |  | NA | NA | No | 0 for drinking water | NA
|  | NA | 0.5 | 0.15 | 0.16 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 20 | NA
|  | NA | 0.5 | NA | 0.061 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 |  |  |  | NA | NA | No | NA | NA
|  | No change | 0.5 | NA | 0.62 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 |  |  |  | NA | NA | No | 50 | NA
|  | No change | 1 | 0.0263 | 0.0699 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 |  |  |  | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.046 | 0.069 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 30 | NA
|  | No change | 1 | 0.087 | 0.065 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% |  |  |  | 5% | NA | No | 122 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | Analyst time and instrument maintenance would increase; 20% cost increase |  |  |  |  | 20 to 50% | NA | No | 125 | NA
|  | No change | 1 | 0.21 | 0.17 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Current reporting limit below proposed DLR | NA | No impact | No impact |  |  |  | No impact | NA | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | 2 | 0.109 | 0.087 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA |  |  |  | NA | NA | No | 25 | NA
|  | No change | 2 | 0.40988 | 0.2817 | As described in 2017 Methods Update Rule | [12 empty] | NA | NA | No | >1000 | NA | bar: Yes to NA
|  | No change | 2 | 0.40988 | 0.2817 | As described in 2017 Methods Update Rule | [12 empty] | NA | NA | No | >1000 | NA | bar: Yes to NA
|  | NA | 20 | 0.396 | 2.25 | 40CFR Part 136, App B | No | 20 | No | additional study | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
|  | No change | NA | 0.09 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | No | 50 | NA
|  | EPA 200.7 | NA | NA | NA | NA | No | 1 | No | NA | NA | Method development. | 12 | 10% | 10% |  |  |  | NA | NA | No | 1 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA |  |  |  | NA | NA | NA | NA | NA
| 20 | No change | 0.1 | 0.06 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 29 | Suite | Yes | 0 | Discounting depends on scope; single element cost = $52 | No | 400 to 600 | NA
|  | No change | 0.1 | 0.074012 | 0.11569 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | NA | NA | Yes | 0 | Pricing depends greatly on other work. Sometimes metals analysis are used as loss leaders. | No | 230 | NA
|  | No change | 0.2 | 0.06 | 0.0217 | as described in 2017 Methods Update Rule | Yes | NA | No | None | NA | None | 0 | 0 | 0 | NA | NA | NA | NA | NA | No | 8 | NA
|  | No change | 0.5 | 0.016 | 0.2 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | 26.79 | Single | NA | NA | NA | NA | 20 | NA
|  | No change | 0.5 | 0.016 | 0.2 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | 26.79 | Single | NA | NA | NA | NA | 20 | NA
|  | NA | 0.5 | 0.03 | 0.12 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | SOP/LIMS update | 1 | NA | NA | NA | NA | No | 0 | NA | No | 0 for drinking water | NA
|  | NA | 0.5 | 0.15 | 0.16 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 20 | NA
|  | NA | 0.5 | NA | 0.061 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | Single | NA | 0 | NA | No | 5 | NA
|  | No change | 0.5 | NA | 0.62 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | 1 | 0.0263 | 0.0699 | As described in 2017 Methods Update Rule | Yes | NA | No | ICP-MS | NA | Staff training | 3 | 0 | 0 | 7 | Suite | NA | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.046 | 0.069 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA, Utility lab | NA, Utility lab | NA | None | NA | No | NA | NA
|  | No change | 1 | 0.087 | 0.065 | As described in method | Yes | NA | No | NA | NA | NA | NA | NA | NA | 30 | Single | 10% | NA | NA | No | 121 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | 0 | Analyst time and instrument maintenance would increase; 20% cost increase |  | 26 | Single | Yes, up to 10% discount; $23.40 | 20% increase; $31.20 | NA | No | 175 | NA
|  | No change | 1 | NA | 0.21 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Current reporting limit below proposed DLR | NA | No impact | No impact | Unknown | Suite | NA | 0 | NA | No | 60 | NA
|  | No change | 2 | 0.109 | 0.087 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | NA | NA | NA | No | 25 | NA
|  | No change | 2 | 0.40988 | 0.2817 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 2 | 0.40988 | 0.2817 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | NA | 20 | 0.396 | 2.25 | 40CFR Part 136, App B | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | NA | 0.09 | 0.1 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | 35 | Single | NA | NA | NA | No | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 50 | NA
|  | NA | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | NA | NA | NA | No | 1 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
SM 3111B | 10 | EPA 200.8 | 500 | 18 | 24 | As described in 2017 Methods Update Rule | No | 75 | Yes | ICP | 150000 | Instrument research & purchase. Installation, Staff training, MDL procedue. | 6 | 10 to 15% | 10% |  |  |  | NA | NA | Yes | 6 to 10 | NA
|  | SM 3113B | NA | NA | NA | NA | No | 20 | No | NA | NA | Method development - staff training | 3 | 0 | 0 |  |  |  | 0 | NA | Yes | 6 per quarter | NA
| 20 | EPA 200.8 | 500 | 18 | 24 | As described in 2017 Methods Update Rule | No | 75 | Yes | ICP | 150000 | Instrument research & purchase. Installation, Staff training, MDL procedue. | 6 | 10 to 15% | 10% | NA | NA | NA | NA | NA | Yes | 6 to 10 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continued Staff training - optimizing instrument | 1 | 0 | 0 | 15 | Single | NA | 0 | NA | No | 6 per quarter | NA

## Mercury
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.8 | 0.2 | No change | 0.2 | 0.0309 | 0.0536 | As described in 2017 Methods Update Rule | No | 0.4 | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 40 | NA
|  | No change | 0.2 | 0.039845 | 0.07905 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 0.2 | 0.039845 | 0.07905 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 0.5 | 0.0624 | 0.0716 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 | 0 | NA | No | 2 | NA
|  | NA | 1 | 0.0321 | 0.056 | 40CFR Part 136, App B | No | 1 | No | additional study | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 1 | 0.1 | 0.3 | As described in 2017 Methods Update Rule | No | 0.3 | Yes | Minimum requirement would be allowing KED (collision cell technology) | New ICPMS and analyst; $150,000 | Trained staff on KED technology or different instrument with those attainable levels | 36 | Analyst time and instrument maintenance would increase; 100%+ cost increase |  | 100 to 200% | NA | Yes | 125 | NA
|  | EPA 245.1 | 1 | 0.15 | 0.43 | As described in 2017 Methods Update Rule | Yes | NA | Yes | Cold Vapor Hg Analyzer | 30000 | Budget for equipment, Staff training, method development, obtain ELAP accreditation. Possible reorganization of scheduling assignments to accommodate up to 1 FTE Senior Chemist dedicated to Hg analysis. | 12 to 18 | 10% | 2% | 5% | Est. $30.28 / sample | No | 60 | Will however work on developing a 200.8 method to try to achieve a 0.2 ug/L DRL
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 77 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
EPA 245.1 | 0.2 | NA | 0.025 | 0.006 | 0.01 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 20 | NA
|  | NA | 0.05 | 0.0005 | 0.006 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | SOP/LIMS update | 6 | NA | NA | NA | NA | No | 0 for drinking water | NA
|  | No change | 0.05 | 0.01 | 0.004 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.05 | 0.0366 | 0.022615 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | 0 | NA | No | 280 | NA
|  | No change | 0.1 | NA | 0.148 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 50 | NA
|  | No change | 0.2 | 0.035 | 0.036 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | No | 10 | NA
|  | No change | 0.2 | 0.058 | 0.055 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 12 | NA
|  | No change | 0.2 | 0.16 | 0.16 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | Current instrument cannot achieve reliably but the instrument is due to be replaced with a newer more sensitive instrument
|  | No change | 0.5 | 0.19 | 0.054 | As described in method | No | 0.5 | Yes | New Hg Analyzer or equivalent | $15,000 to $30,000 | Method development, Staff Training - New equipment | 12 to 24 | 100% | 25% | 5% | NA | Yes | 27 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 50 | NA
|  | None, subcontract | NA | NA | NA | NA | No | 1 | No | NA | NA | Method development | 12 | 10% | 10% | NA | NA | No | 10 | NA
SM 3120B | 0.2 | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | Continued Staff training - optimzing instrument | 1 | 0 | 0 | 0 | NA | No | 6 per quarter | NA

## Thallium
Current Analytical Method | Draft DLR (ug/L) | Method lab would change to if current method could not be used to meet future DLR | Lowest Calibration Point (ug/L) | MDL - blank (ug/L) | MDL - spiked (ug/L) | MDL Determination Procedure | Can you achieve the DLR with +/- 30% recovery range at the DLR? | If NO, what DLR can you achieve with +/- 30% recovery range? (ug/L) | Need new equipment to achieve DLR? | Specific Equipment Needed | Approximate Equipment Cost | What specific actions are needed to build capacity to achieve this DLR? | Realistic timeline to complete the actions (months) | Capital Cost Impacts: Approximate Cost Increase (%) | Operating Cost Impacts: Approximate Cost Increase (%) | Expected Increase in Analysis Cost (%) | Other Pricing Comments | At draft DLR, would you lose ELAP accreditation for this analyte? | Approximate samples analyzed per month | Comments
EPA 200.8 | 0.5 | No change | 0.05 | 0.08 | 0.05 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.1 | 0.00318 | 0.020104 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | 15% | NA | No | 130 | NA
|  | No change | 0.2 | 0.06 | 0.04 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 | NA | NA | NA | 8 | NA
|  | No change | 0.5 | 0.003 | 0.06 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 3 | NA
|  | No change | 0.5 | 0.003 | 0.06 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 3 | NA
|  | NA | 0.5 | 0.004 | 0.05 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | MDL study at lower spiking concentration, SOP/LIMS update, possibly method development | 6 | NA | NA | NA | NA | No | 0 for drinking water | NA
|  | NA | 0.5 | 0.068 | 0.084 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 20 | NA
|  | NA | 0.5 | NA | 0.031 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | NA | NA
|  | No change | 0.5 | NA | 0.11 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 50 | NA
|  | No change | 1 | 0.013 | 0.113 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% | 5% | NA | NA | NA | NA
|  | No change | 1 | 0.014 | 0.014 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 30 | NA
|  | No change | 1 | 0.0266 | 0.1701 | As described in 2017 Methods Update Rule | No | 2 | Yes | ICP-MS with DRC | 250000 | staff training, method development | 6 | 200% | NA | NA | NA | Yes | 25 | DRC is currently not acceptable for drinking water accreditation.
|  | No change | 1 | 0.058 | 0.062 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | No | 11 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | No | 0.5 | Yes | Minimum requirement would be allowing KED (collision cell technology) | New ICPMS and analyst; $150,000 | Trained staff on KED technology or different instrument with those attainable levels | 36 | Analyst time and instrument maintenance would increase; 100%+ cost increase |  | 100 to 200% | NA | Yes | 125 | NA
|  | No change | 1 | 0.1364 | 0.2087 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.32291 | 0.052049 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.43 | 0.31 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development | 3 to 6 | 2% | No impact | 2% | Primary factors include staff testing existing instrument capabilities, potentially investing in new custom calibration standards, etc. to meet proposed DLR | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | NA | 0.02 | 0.07 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | NA | NA
|  | No change | NA | NA | NA | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | None, subcontract | NA | NA | NA | NA | No | 0.8 | No | NA | NA | Method development. | 6 | 5% | 5% | NA | NA | No | 1 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
| 0.8 | No change | 0.05 | 0.08 | 0.05 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | No foreseeable addt'l actions needed | NA | 0 | 0 | 0 | NA | No | 400 to 600 | NA
|  | No change | 0.1 | 0.00318 | 0.020104 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | None | NA | NA | NA | 15% | NA | No | 130 | NA
|  | No change | 0.2 | 0.06 | 0.04 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | 0 | None | NA | 0 | 0 | NA | NA | NA | 8 | NA
|  | No change | 0.5 | 0.003 | 0.06 | EPA 821-R-16-006, December 2016 | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 3 | NA
|  | NA | 0.5 | 0.004 | 0.05 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | MDL study at lower spiking concentration, SOP/LIMS update, possibly method development | 6 | NA | NA | NA | NA | No | 0 for drinking water | NA
|  | NA | 0.5 | NA | 0.031 | As described in method | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | NA | NA
|  | No change | 0.5 | NA | 0.11 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 50 | NA
|  | NA | 1 | 0.0041 | 0.0809 | 40CFR Part 136, App B | No | 1 | No | additional study | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | 1 | 0.013 | 0.113 | As described in method | No | 1 | No | NA | NA | Method development, Staff Training | 3 | 2% | 5% | 5% | NA | No | 18 | NA
|  | No change | 1 | 0.014 | 0.014 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 30 | NA
|  | No change | 1 | 0.0266 | 0.1701 | As described in 2017 Methods Update Rule | No | 2 | Yes | ICP-MS with DRC | 250000 | staff training, method development | 6 | 200% | NA | NA | NA | Yes | 25 | DRC is currently not acceptable for drinking water accreditation.
|  | No change | 1 | 0.058 | 0.062 | As described in method (Method references Code of Federal Regulations 40, Ch. 1, Pt. 136 Appendix B) | Yes | NA | No | NA | NA | Update SOP, Purchase custom standard | 1 | NA | NA | NA | NA | No | 11 | NA
|  | No change | 1 | 0.1 | 0.5 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | Analyst time and instrument maintenance would increase; 20% cost increase |  | 20 to 50% | NA | No | 125 | NA
|  | No change | 1 | 0.1364 | 0.2087 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development, staff training | 12 | 0 | 0 | 0 | NA | No | 2 | NA
|  | No change | 1 | 0.32291 | 0.052049 | As described in 2017 Methods Update Rule | No | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | >1000 | NA
|  | No change | 1 | 0.43 | 0.31 | As described in 2017 Methods Update Rule | Yes | NA | No | NA | NA | Method development | 3 to 6 | 2% | No impact | 2% | Primary factors include staff testing existing instrument capabilities, potentially investing in new custom calibration standards, etc. to meet proposed DLR | No | 60 | For DLR elements <1 ug/L while achieving +/-30% recovery is doable, it will cause more recalibration/reanlaysis do to a higher rate for failures. Recommend allowance of +/- 50% at DLR level.
|  | No change | NA | 0.02 | 0.07 | As described in 2017 Methods Update Rule | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
|  | No change | NA | NA | NA | NA | Yes | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | No | 50 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | No | 59 | NA
|  | None, subcontract | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | 0 | 0 | NA | NA | No | 1 | NA
|  | No change | NA | NA | NA | NA | Yes | NA | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA
EPA 200.9 | 0.5 | EPA 200.8 | NA | NA | 0.44 | As described in method | No | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | NA | 2 | NA

## Miscellaneous
Please explain cost increases for the draft DLRs if they are higher or lower than an increase of… |  |  |  | Number of Employees
43% for EPA Method 200.7 | 20% for EPA Method 200.8 | 45% for EPA Method 245.1 | 40% for SM 3113B
Training and Method Develop costs as well as Capital Expenditures | Training and Method Develop costs as well as Capital Expenditures | Training and Method Develop costs as well as Capital Expenditures | Training and Method Develop costs as well as Capital Expenditures | 65
Does not exceed maximum analytical cost increase | Does not exceed maximum analytical cost increase | Does not exceed maximum analytical cost increase | Does not exceed maximum analytical cost increase | 37
If our ICP OES can't meet Fe DLR we will to purchase a new one | the biggest cost in this area is method development | Current technology cant meet new DLR new instrument would need to be purchased | NA | 21
Unknown, this is market driven | Unknown, this is market driven | NA | NA | 130
NA | Metals can be determined at a lower level if drinking water will allow collision cell technology (KED) for ICPMS *note: wastewater already allows KED | NA | NA | 2
NA | Considerable method development, switch from Argon plasma to Helium plasma (~20x cost increase), improve the building's electrical infrastructure, possibly buy a new instrument, etc. | NA | NA | 16
NA | If new instrument purchase is needed to address lower DLR for arsenic, then increase in capital cost is expected, which will impact the individual cost of analyses. | NA | NA | 3 staff currently performing metals analysis
NA | NA | We estimate 45 to 50% cost increase per analysis based on increased staff time for senior analyst, amortization of capital equipment purchase, instrument service contract, and increased reagent and other consumable costs. Hg would become a standalone analysis. without benefit of economies of scale currently realized by including with other Method 200.8 analyses. | NA | 33 full-time, 5 part-time
NA | NA | NA | To change analytical method from SM 3113B to EPA Method 200.8, will change from individual analysis to suite analysis so 130% increase will be divided by 2 for the lead and arsenic tests or 65% for each | 6
NA | NA | NA | NA | 30 (1 for EPA 200.8)
NA | NA | NA | NA | 5
NA | NA | NA | NA | Lab 24/Metal 8
NA | NA | NA | NA | 8
NA | NA | NA | NA | 75
NA | NA | NA | NA | 3
NA | NA | NA | NA | 5
NA | NA | NA | NA | 3
NA | NA | NA | NA | 6
NA | NA | NA | NA | 43
NA | NA | NA | NA | 1 (Metals only)
NA | NA | NA | NA | EPA 200.8 -three employees; EPA 200.7 - three employees; EPA 245.1 - two employees
NA | NA | NA | NA | 9
NA | NA | NA | NA | Metals + Extr 10 employees